# APPENDIX  I                                                          XYZ/CQA/SOP-040/FR-01

# LOG BOOK FOR FTIR

# INSTRUMENT ID ------------------------------------

# LOCATION         -------------------------------------

| # Sr.No | # Date | # Name of Product | # Batch No/ A.R No | # Done by | # Checked by | # Remarks |
| --- | --- | --- | --- | --- | --- | --- |

XYZ/CQA/SOP-040/FR-02

APPENDIX II

CALIBRATION DATA SHEET              INSCAL------------------------

INSTRUMENT ID              	              : --------------------------

LOCATION		         	              : --------------------------

FREQUENCY OF CALIBRATION       : QUARTERLY

DATE OF CALIBRATION	               :--------------------------

NEXT DUE DATE OF CALIBRATION :--------------------------

(A) WAVENUMBER CALIBRATION

| S. No. | Standard Wave Number.(cm-1) | Observed Value of Wave Number (cm-1) | Tolerance (cm-1) |
| --- | --- | --- | --- |
| 1. | 3060 |  | 3059 - 3061 |
| 2. | 2849.5 |  | 2848.5 - 2850.5 |
| 3. | 1942.9 |  | 1941.9 - 1943.5 |
| 4. | 1601.2 |  | 1600.2 - 1602.2 |
| 5. | 1583 |  | 1582 - 1584 |
| 6. | 1154.5 |  | 1153.5 - 1155.5 |
| 7. | 1028.3 |  | 1027.3 - 1029.3 |

Acceptance criteria  :                       1 cm-1  of the Standard wave number.

Remarks :                                         Complies /  Does not Comply

Done by         :				Checked By	:		Approved By   :

Date		:			Date		:		Date		:

(B)   RESOLUTION PERFORMANCE :

1.0 The difference of transmittance between the depth of the trough from the minimum

Transmission at 2870 cm-1 to the maximum at 2849.5 ± 1.0 cm-1.

| Minimum Transmittance at 2870 cm-1 | Maximum Transmittance at 2849.5 cm-1 | Difference in Transmittance | Acceptance Criteria |
| --- | --- | --- | --- |
|  |  |  | The difference should be greater than 18%Transmittance |

2.0The difference of transmittance between the depth of the trough from the maximum

Transmission at 1589 cm-1 to the minimum at 1583 ± 1.0 cm-1

| Minimum Transmittance at 1589 cm-1 | Maximum Transmittance at 1583 cm-1 | Difference in Transmittance | Acceptance Criteria |
| --- | --- | --- | --- |
|  |  |  | The difference should be greater than 10%Transmittance |

Remarks :                                          Complies /  Does not Comply

Done by         :				Checked By	:		Approved By   :

Date		:			Date		:		Date		: