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INSTRUMENT ID ------------------------------------

LOCATION         -------------------------------------   

	Sr.No
	Date
	Name of Product
	Batch No/ A.R No
	Done by
	Checked by
	Remarks


	
	
	
	
	
	
	


	
	
	
	
	
	
	


	
	
	
	
	
	
	


	
	
	
	
	
	
	


	
	
	
	
	
	
	


	

	
	
	
	
	
	

	

	
	
	
	
	
	

	

	
	
	
	
	
	

	

	
	
	
	
	
	

	

	
	
	
	
	
	




	
	Code
	:
	xyz/CQA/SOP-040

	
	Issue No.
	:
	01

	
	Effective Date
	:
	27-Jan-2010

	  OPERATION  AND CALIBRATION OF FTIR(JASCO-FT/IR 410)
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                                                                                         XYZ/CQA/SOP-040/FR-02
	APPENDIX II

[bookmark: _Toc120003142]	CALIBRATION DATA SHEET              INSCAL------------------------

INSTRUMENT ID              	              : --------------------------
LOCATION		         	              : --------------------------    
FREQUENCY OF CALIBRATION       : QUARTERLY
DATE OF CALIBRATION	               :--------------------------				
NEXT DUE DATE OF CALIBRATION :--------------------------

(A) WAVENUMBER CALIBRATION


	S. No.
	Standard Wave Number.(cm-1)
	Observed Value of Wave Number (cm-1)
	Tolerance
(cm-1)

	1.
	3060
	
	3059 - 3061

	2.
	2849.5
	
	2848.5 - 2850.5

	3.
	1942.9
	
	1941.9 - 1943.5

	4.
	1601.2
	
	1600.2 - 1602.2

	5.
	1583
	
	1582 - 1584

	6.
	1154.5
	
	1153.5 - 1155.5

	7.
	1028.3
	
	1027.3 - 1029.3














Acceptance criteria  :                       1 cm-1  of the Standard wave number.

Remarks :                                         Complies /  Does not Comply

Done by         :				Checked By	:		Approved By   :

Date		:			Date		:		Date		:	

 



 (B)   RESOLUTION PERFORMANCE :


    1.0 The difference of transmittance between the depth of the trough from the minimum    
         Transmission at 2870 cm-1 to the maximum at 2849.5 ± 1.0 cm-1.

	Minimum Transmittance at 2870 cm-1                                          
	Maximum Transmittance at 2849.5 cm-1 
	Difference in Transmittance
	Acceptance Criteria

	
	
	
	The difference should be greater than 18%Transmittance








             
  2.0The difference of transmittance between the depth of the trough from the maximum 
 Transmission at 1589 cm-1 to the minimum at 1583 ± 1.0 cm-1


	Minimum Transmittance at 1589 cm-1                                          
	Maximum Transmittance at 1583 cm-1 
	Difference in Transmittance
	Acceptance Criteria

	 
	
	
	The difference should be greater than 10%Transmittance

















Remarks :                                          Complies /  Does not Comply

Done by         :				Checked By	:		Approved By   :

Date		:			Date		:		Date		:	


