APPENDIX  I                         XYZ/CQA/SOP-076/FR-01

CALIBRATION DATA SHEET           INSCAL…………………….

- Micropipette id:……………………………

- Location          :…………………………….

- Frequency       : Half Yearly

- Date of Calibration :-………………………

- Next due date of Calibration……………..

| Range of Micropipette | Wt of water In mg |  | Average Weight | Standard Deviation | % CV | V = Avg wt 0.99602 | A% | Remarks |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Low | I |  |  |  |  |  |  |  |
|  | II |  |  |  |  |  |  |  |
|  | III |  |  |  |  |  |  |  |
|  | IV |  |  |  |  |  |  |  |
|  | V |  |  |  |  |  |  |  |
| Mid | I |  |  |  |  |  |  |  |
|  | II |  |  |  |  |  |  |  |
|  | III |  |  |  |  |  |  |  |
|  | IV |  |  |  |  |  |  |  |
|  | V |  |  |  |  |  |  |  |
| High | I |  |  |  |  |  |  |  |
|  | II |  |  |  |  |  |  |  |
|  | III |  |  |  |  |  |  |  |
|  | IV |  |  |  |  |  |  |  |
|  | V |  |  |  |  |  |  |  |

Calculation of  A %

A% 	 =	 V   -   VSET	X 100

VSET

For Low Range      Vset  =

A%                      ___________    X     100

For  Mid Range      Vset   =

A%                      ___________    X     100

Page 2 of 2

For  High Range    Vset   =

A%                      ___________    X     100

Acceptance Criteria

| Range of Micropipette | A% | A% in µl | % CV |
| --- | --- | --- | --- |
| 0.1-1 µl | ±5.0 | 0.05 | 2.5 |
| 0.5-10 µl | ±1.0 | 0.2 | 0.5 |
| 2-20 µl | ±1.0 | 0.2 | 0.5 |
| 5-50 µl | ±1.0 | 0.5 | 0.4 |
| 10-100 µl | ±0.8 | 0.8 | 0.3 |
| 20-200 µl | ±0.8 | 1.6 | 0.3 |
| 25-250 µl | ±0.8 | 2.0 | 0.3 |
| 100-1000 µl | ±0.8 | 8.0 | 0.3 |
| 500-5000 µl | ±0.8 | 40 | 0.3 |

Remarks: The above parameter used for calibration as per predefined acceptance criteria

complies /  Does not comply hence the micropipette  is suitable /Not suitable for

routine analysis.

Done By:				Checked By:				Approved By:

Date:					Date:	                                                Date: