APPENDIX   I	                   XYZ/CQA/SOP-051/FR-01

LOG  BOOK  FOR  GC

Instrument Id            : ---------------------------

Location                   : ----------------------------

| Sr.No | Date | Name o f Product | Batch No /AR.No | Done By | Checked by | Remarks |
| --- | --- | --- | --- | --- | --- | --- |

APPENDIX II                             XYZ/CQA/SOP-051/FR-02

CALIBRATION DATA SHEET                INSCAL------------------------

Page  1 of 9

- Instrument ID            : ------------------------

- Location                     :------------------------

3.  Calibration frequency : Half yearly

4.  Date of Calibration      :------------------------

5.  Next due date of calibration:------------------

5.  Flow meter ID             :------------------------

6.  Validity for calibration :-------------------------

7.0 Calibration for Gas Flow Rate using Digital Flow meter

7.1 Observations

| Sr.No | Set Flow Rate | Observed Flow rate |  |  |  |  | Average Flow rate | Acceptance Criteria |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | I | II | III | IV | V |  |  |
| 1.0 | 5ml/min |  |  |  |  |  |  | ± 5% of set flow rate |
| 2.0 | 10ml/min |  |  |  |  |  |  |  |
| 3.0 | 15ml/min |  |  |  |  |  |  |  |
| RSD |  |  |  |  |  |  |  | NMT2% |

Remarks :                                          Complies / Does not comply

Done by  :                                          Checked by :                                           Approved  by :

Date       ;                                           Date           :                                            Date               :

- Calibration of Gas Flow Rate using FID detector                             Page 2 of 9

8.1	Observations :

| Sr.No | Set Flow Rate | Observed Flow rate | Flow Accuracy | Acceptance Criteria |
| --- | --- | --- | --- | --- |
| 1.0 | 5ml/min |  |  | ± 5% of set flow rate |
| 2.0 | 10ml/min |  |  |  |
| 3.0 | 15ml/min |  |  |  |

Calculation of Flow Accuracy ;

Flow accuracy       :      Observed flow rate           X     100

Set Flow rate

Remarks :                                          Complies / Does not comply

Done by  :                                          Checked by :                                           Approved  by :

Date       ;                                           Date           :                                            Date               :

9.0	Calibration for Temperature :                                                          Page 3 of 9

  - In Column Oven

9.1.1    Observations :

| Sr.No | Set Temperature | Observed Temperature | Acceptance criteria |
| --- | --- | --- | --- |
| 1 | 40°C |  | ± 2°C of set temperature |
| 2 | 100°C |  |  |
| 3 | 150°C |  |  |
| 4 | 200°C |  |  |
| 5 | 240°C |  |  |

Remarks :                                          Complies / Does not comply

Done by  :                                          Checked by :                                           Approved  by :

Date       ;                                           Date           :                                            Date

9.2    In Injection Port (Calibration for Temperature)                                        Page 4 of 9

9.2.1 Observations

| Sr.No | Set Temperature | Observed Temperature | Acceptance criteria |
| --- | --- | --- | --- |
| 1 | 140°C |  | ± 10°C of set temperature |
| 2 | 200°C |  |  |
| 3 | 280°C |  |  |

Remarks :                                           Complies / Does not comply

Done by  :                                          Checked by :                                           Approved  by :

Date       ;                                           Date           :                                            Date

9.3  	In Detector (Calibration for temperature)                                            Page 5 of 9

9.3.1	Observations

| Sr.No | Set Temperature | Observed Temperature | Acceptance criteria |
| --- | --- | --- | --- |
| 1 | 180°C |  | ± 10°C of set temperature |
| 2 | 250°C |  |  |
| 3 | 300°C |  |  |

Remarks :                                           Complies / Does not comply

Done by  :                                          Checked by :                                           Approved  by :

Date       ;                                           Date           :                                            Date

  - In Head space Incubator ( Calibration for temperature)                  Page 6 of 9

9.4.1 	Observations

| Sr.No | Set Temperature | Observed Temperature | Acceptance criteria |
| --- | --- | --- | --- |
| 1 | 80°C |  | ± 2°C of set temperature |
| 2 | 85°C |  |  |
| 3 | 90°C |  |  |

Remarks :                                           Complies / Does not comply

Done by  :                                          Checked by :                                           Approved  by :

Date       ;                                           Date           :                                            Date

10.0	Linearity of Detector Response                                                         Page 7 of 9

10.1	Preparation of Stock Solution

Weigh accurately--------------mg of Ethyl Acetate,------------mg of isopropyl alcohol,------------- mg of Methylene Chloride in ---------------ml of volumetric flask containing ----------------ml

of   purified  water.

Dilute --------ml of Stock solution to ---------- ml with purified water.(100ppm)

Dilute---------ml of Stock solution to ----------ml with purified water (50ppm)

Dilute --------ml of  Stock solution to ---------ml  with purified water (25ppm)

| Concentration | Injection volume | Peak Area response |  |  |  |  | Average | %RSD |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | I | II | III | IV | V |  |  |
| 25ppm | 2µl |  |  |  |  |  |  |  |
| 50ppm | 2µl |  |  |  |  |  |  |  |
| 100ppm | 2µl |  |  |  |  |  |  |  |
| Slope |  |  |  |  |  |  |  |  |
| Intercept |  |  |  |  |  |  |  |  |
| Correlation Coefficient |  |  |  |  |  |  |  |  |

| Concentration | Retention time |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | I | II | III | IV | V |
| 25ppm |  |  |  |  |  |
| 50ppm |  |  |  |  |  |
| 100ppm |  |  |  |  |  |

Acceptance criteria : The % RSD of area of six injection of 25ppm should be less than 10%

and retention time should be less than 2%.

The correlation coefficient should not be less than 0.98.

Remarks :                                           Complies / Does not comply

Done by  :                                          Checked by :                                           Approved  by :

Date       ;                                           Date           :                                            Date

- Detector ‘ A ‘ sensitivity  (FID packed Column):                               Page 8 of 9

Calculation  of  Detector  sensitivity :

Sensitivity (In coulombs per gram )  = ( H  X WH )

V

where	  =  (Ampere gain range 20) x (Integrator Atten) x (signal ‘ mv’ )

[ Ampere gain range 20 =   7 x 10-12 Amp / mv,

Integrator Atten           =    64  ]

WH  = (Width at half  height in seconds)

( Note: width at half height in Minutes x 60  = WH in  seconds)

Remarks :                                           Complies / Does not comply

Done by  :                                          Checked by :                                           Approved  by :

Date       ;                                           Date           :                                            Date

12.0	Detector ‘ A ‘ sensitivity   (FID Capillary Column):                          Page 9of 9

Sensitivity (In coulombs per gram)     =  H X WH       X       1 + SP

V                        C

H     =  (Ampere gain range 1) x (Integrator Atten) x (signal ‘ mv’ )

[ Ampere gain range 1  =   3.33 x 10-13 Amp / mv ,

Integrator Atten           = 8]

WH  = (Width at half height in seconds)

(Note: width at half height in Minutes x 60  = WH in seconds)

V     = sample weight C8 in 0.5µl

V     =  ( 2.965 x 10 –6 ) Grams carbon

SP   = Split flow in ml / min  = 50 ml / min

C     = Column flow in ml / min  = 1 ml / min

Acceptance Criteria:

Detector ‘A’ sensitivity FID Capillary Column): should be more than 0.013c/g

Remarks :                                           Complies / Does not comply

Done by  :                                          Checked by :                                           Approved  by :

Date       ;                                           Date           :                                            Date

APPENDIX  III                                                       XYZ/CQA/SOP-051/FR-03

GC CALIBRATION SUMMARY SHEET                                   INSCAL -----------------------------

Instrument id            :  -----------------------------------

Location                    : -------------------------------------

| Test Name | Observation |  |  |  |  |  |  |  |  |  |  |  |  | Acceptance Criteria | Status |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Calibration of Gas flow rate Using flow meter | Sr. No. | Set Flow rate ml/min |  | Observed flow rate (ml/min) |  |  |  |  |  |  |  | Average flow rate | Mean %RSD | ± 5% of set flow rate RSD NMT2% | Complies / Does not comply |
|  |  |  |  | I |  |  | II | III | IV | V |  |  |  |  |  |
|  | 1 | 5 |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 2 | 10 |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 3 | 15 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Calibration of Gas Flow Rate using FID detector | Sr.No |  | Set Flow rate |  |  |  | Observed flow rate |  |  |  |  |  |  | ± 5% of set flow rate | Complies / Does not comply |
|  | 1 |  | 5ml/min |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 2 |  | 10ml/min |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 3 |  | 15ml/min |  |  |  |  |  |  |  |  |  |  |  |  |
| Calibration for Temperature In Column Oven | Sr.No |  | Set temp. |  |  |  | Observed temp. |  |  |  |  |  |  | ± 2°C of set temperature | Complies / Does not comply |
|  | 1 |  | 40°C |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 2 |  | 100°C |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 3 |  | 150°C |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 4 |  | 200°C |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 5 |  | 240°C |  |  |  |  |  |  |  |  |  |  |  |  |
| Calibration of Temperature (In Injection Port) | Sr.No |  | Set Temp. |  |  | Observed Temp |  |  |  |  |  |  |  | ± 10°C of set temperature | Complies / Does not comply |
|  | 1 |  | 140 °C |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 2. |  | 200°C |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 3 |  | 280°C |  |  |  |  |  |  |  |  |  |  |  |  |
| In Detector | Sr.No |  | Set Temp. |  |  | Observed Temp |  |  |  |  |  |  |  | ± 10°C of set temperature | Complies / Does not comply |
|  | 1 |  | 180 °C |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 2 |  | 250 °C |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 3 |  | 300 °C |  |  |  |  |  |  |  |  |  |  |  |  |
| In Head space Incubator | Sr.No |  | Set Temp. |  |  | Observed Temp |  |  |  |  |  |  |  | ± 2°C of set temperature | Complies / Does not comply |
|  | 1 |  | 80 °C |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 2 |  | 85 °C |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 3 |  | 90 °C |  |  |  |  |  |  |  |  |  |  |  |  |
| Linearity of Detector response (Correlation Coefficient and %RSD) | Solvent Name |  |  |  | Correlation Coefficient |  |  |  |  |  | %RSD |  |  | Correlation Coefficient should not be less than 0.98 %RSD should be less than 10 %. | Complies / Does not comply |
|  | IPA |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | Dichloromethane |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | Ethyl Acetate |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| Detector ‘ A ‘ sensitivity (FID packed Column): |  | Detector ‘A’ sensitivity should be more than 0.015c/g | Complies / Does not comply |
| --- | --- | --- | --- |
| Detector ‘ A ‘ sensitivity (FID Capillary Column): |  | Detector ‘A’ sensitivity should be more than 0.013c/g | Complies / Does not comply |

Remarks:    All parameters used for the calibration as per predefined acceptance criteria complies / Does not comply hence the instrument is suitable / Not suitable  for  routine analysis

Done by :                                                                              Checked by :                                                           Approved by  :

Date      ;                                                                               Date           :                                                            Date              :