APPENDIX I                                                         ZYX /CQA/SOP-039/FR-01

LOG BOOK OF TOC ANALYSER

Instrument ID       :	-------------------------------

Location	       :	-------------------------------

| Sr. No | Date | Name of Sample | Batch No / A.R. No | Done by | Checked by | Remarks |
| --- | --- | --- | --- | --- | --- | --- |

ZYX/CQA/SOP-039/FR-02                                                            APPENDIX II

CALIBRATION BY POTASSIUM HYDROGEN PHTHALATE

(Note: Calibration shall be done on the day of sample analysis)

1.0	Instrument ID .			: -------------------------

2.0	Location				: --------------------------

3.0 	Date of Calibration			: ---------------------------

4.0	Preparation of Standard solution

Wt of KHP (Batch/ Lot No------------------Make-------------). -------------------gm

dissolved in -------------ml water (1000 ppm)

Dilute -----------ml of solution(1000ppm) to --------------ml of purified water.(10ppm)

Dilute -----------ml of solution (10ppm)    to --------------ml of purified water (250ppb)

Dilute -----------ml of solution (10ppm)    to-------------- ml of purified water (500ppb)

Dilute -----------ml of solution (10ppm)    to --------------ml of purified water(750ppb)

Dilute -----------ml of solution (10ppm)    to --------------ml of purified water(1000ppb)

Dilute -----------ml of solution (10ppm)    to --------------ml of purified water(2000ppb)

Dilute -----------ml of solution (10ppm)    to --------------ml of purified water(5000ppb)

Observations :

| Sr.No | Concentration | Area Count |  |  |  |  | Average Area Count | SD | RSD |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | I | II | III | IV | V |  |  |  |
| 1. | 10ppm |  |  |  |  |  |  |  |  |
| 2. | 0ppb |  |  |  |  |  |  |  |  |
| 3. | 250ppb |  |  |  |  |  |  |  |  |
| 4. | 500ppb |  |  |  |  |  |  |  |  |
| 5. | 750ppb |  |  |  |  |  |  |  |  |
| 6. | 1000ppb |  |  |  |  |  |  |  |  |
| 7. | 2000ppb |  |  |  |  |  |  |  |  |
| 8. | 5000ppb |  |  |  |  |  |  |  |  |
| Slope |  |  |  |  |  |  |  |  |  |
| Intercept |  |  |  |  |  |  |  |  |  |
| Correlation coefficient |  |  |  |  |  |  |  |  |  |

Acceptance criteria  : Linearity  : Correlation coefficient  should not be less than 0.98

Standard Deviation:  NMT 1.0

RSD (C.V.)             : NMT 5.0%

Remarks :				Complies / Does not Comply

Done by  :				Checked By	:			Authorised By	:

Date        :				Date	 :	           		Date                 :

APPENDIX III                             ZYX/CQA/SOP-039/FR-03

CALIBRATION BY SUCROSE (Alternate Method)

(Note : Calibration shall be done on the day of sample analysis)

1.0 Instrument ID..		: -------------------------

2.0 Location			: -------------------------

3.0 Date of Calibration	 :-------------------------

4.0Preparation of Standard solution

Wt of Sucrose (Batch/ Lot No------------------Make-------------)------------------.gm

dissolved in -------------ml water  (50 ppm)

Dilute-------- ml of solution (50ppm) to---------------ml Purified water    (250ppb)

Dilute-------- ml of solution (50ppm) to---------------ml Purified water    (500ppb)

Dilute-------- ml of solution (50ppm) to---------------ml Purified water    (1000ppb)

| Sr.No | Concentration | Area Count |  |  |  |  | Average Area Count | SD | RSD |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | I | II | III | IV | V |  |  |  |
| 1. | 0ppb |  |  |  |  |  |  |  |  |
| 2. | 250ppb |  |  |  |  |  |  |  |  |
| 3. | 500ppb |  |  |  |  |  |  |  |  |
| 4. | 1000ppb |  |  |  |  |  |  |  |  |
| Slope |  |  |  |  |  |  |  |  |  |
| Intercept |  |  |  |  |  |  |  |  |  |
| Correlation coefficient |  |  |  |  |  |  |  |  |  |

Acceptance criteria  : Linearity  : Correlation coefficient  should not be less than 0.98

Standard Deviation :  NMT 1.0

RSD                         : NMT  5.0%

Remarks :				Complies / Does not Comply

Done by  :				Checked By	:			Authorised By	:

Date        :				Date		::	           		Date		:

APPENDIX IV                         ZYX/CQA/SOP-039/FR-04

SYSTEM SUITABILITY BY 1,4 BENZOQUINONE (WEEKLY)

1.0 	Instrument ID                    		 :-------------------				:

2.0 	Location		            	:---------------------

3.0 	Date of System Suitability		:----------------------

4.0	Preparation of Standard solution

4.1	Wt of 1,4 Benzoquinone (Batch/ Lot No------------------Make-------------)------------------.gm

dissolved in -------------ml water  (50 ppm)

Dilute-------- ml of solution (50ppm) to---------------ml Purified water    (500ppb)

  - Observations:

Mean Area of Sucrose----------------

Mean area of 1.4 Benzoquinone----------------.

Mean area  of  pure water------------------

4.3	Calculation of Response efficiency (% recovery)

%	R =  (Rss - Rw)

----------------    X 100

(Rs-Rw)

Where Rss  = Mean area counts of 1,4 Benzoquinone solution

Rs   = Mean area counts of sucrose

Rw  = Mean area counts of pure water

Acceptance criteria : The system is said to be passing the test if the % recovery in     between 85% to 115%.

Remarks :			Complies / Does not comply

Done by ;			Checked by	;			Authorized by	:

Date:				Date		 :			Date

APPENDIX  V                                                          ZYX/CQA/SOP-039/FR-05

INSCAL--------------------

CALIBRATION  BY POTASSIUM HYDROGEN PHTHALATE, OR SUCROSE, 1.4 BENZOQUINONE

(HALF YEARLY RECORD)

1.0 Instrument ID                                   :------------------------------

2.0 Location                                           :------------------------------

3.0 Date of Calibration                           :------------------------------

4.0 Next Due date ofCalibration             :------------------------------

5.0 By KHP

5.1 Preparation of Standard solution

5.2 Wt of KHP (Batch/ Lot No------------------Make-------------)------------------.gm ,  dissolved in -------------ml water     (1000ppm)

Dilute -----------ml of solution(1000ppm) to --------------ml of purified water.(10ppm)

Dilute -----------ml of solution (10ppm)    to --------------ml of purified water (250ppb)

Dilute -----------ml of solution (10ppm)    to-------------- ml of purified water (500ppb)

Dilute -----------ml of solution (10ppm)    to --------------ml of purified water(750ppb)

Dilute -----------ml of solution (10ppm)    to --------------ml of purified water(1000ppb)

Dilute -----------ml of solution (10ppm)    to --------------ml of purified water(2000ppb)

Dilute -----------ml of solution (10ppm)    to --------------ml of purified water(5000ppb)

Observations :

| Sr.No | Concentration | Area Count |  |  |  |  | Average Area Count | SD | RSD | Acceptance criteria |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | I | II | III | IV | V |  |  |  | Correlation coefficient should not be less than 0.98 SD : NMT 1.0 RSD : NMT 5.0% |
| 1. | 10ppm |  |  |  |  |  |  |  |  |  |
| 2. | 0ppb |  |  |  |  |  |  |  |  |  |
| 3. | 250ppb |  |  |  |  |  |  |  |  |  |
| 4. | 500ppb |  |  |  |  |  |  |  |  |  |
| 5. | 750ppb |  |  |  |  |  |  |  |  |  |
| 6. | 1000ppb |  |  |  |  |  |  |  |  |  |
| 7. | 2000ppb |  |  |  |  |  |  |  |  |  |
| 8. | 5000ppb |  |  |  |  |  |  |  |  |  |
| Slope |  |  |  |  |  |  |  |  |  |  |
| Intercept |  |  |  |  |  |  |  |  |  |  |
| Correlation coefficient |  |  |  |  |  |  |  |  |  |  |

6.0 By Sucrose

6.1 Preparation of Standard solution

Wt of Sucrose (Batch/ Lot No------------------Make-------------)------------------.gm

dissolved in -------------ml water  (50 ppm)

Dilute-------- ml of solution (50ppm) to---------------ml Purified water    (250ppb)

Dilute-------- ml of solution (50ppm) to---------------ml Purified water    (500ppb)

Dilute-------- ml of solution (50ppm) to---------------ml Purified water    (1000ppb)

| Sr.No | Concentration | Area Count |  |  |  |  | Average Area Count | SD | RSD |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | I | II | III | IV | V |  |  |  |
| 1 | 0ppb |  |  |  |  |  |  |  |  |
| 2. | 250ppb |  |  |  |  |  |  |  |  |
| 3. | 500ppb |  |  |  |  |  |  |  |  |
| 4. | 1000ppb |  |  |  |  |  |  |  |  |
| Slope |  |  |  |  |  |  |  |  |  |
| Intercept |  |  |  |  |  |  |  |  |  |
| Correlation coefficient |  |  |  |  |  |  |  |  |  |

Acceptance criteria  : Linearity  : Correlation coefficient  should not be less than 0.98

Standard Deviation :  NMT 1.0

RSD                         : NMT  5.0%

7.0 By 1,4 Benzoquinone

Preparation of Standard solution

7.1Wt of 1’4 Benzoquinone (Batch/ Lot No------------------Make-------------)------------------.gm

dissolved in -------------ml water  (50 ppm)

Dilute-------- ml of solution (50ppm) to---------------ml Purified water    (500ppb)

7.2	Observations:

Mean Area of Sucrose----------------

Mean area of 1.4 Benzoquinone----------------.

Mean area  of  pure water------------------

7.3	Calculation of Response efficiency (% recovery)

%	R =  (Rss - Rw)

----------------    X 100

(Rs-Rw)

Where   Rss  = Mean area counts of 1,4 Benzoquinone solution

Rs   = Mean area counts of sucrose

Rw  = Mean area counts of pure water

Acceptance criteria : The system is said to be passing the test if the % recovery in   between 85% to 115%.

Remarks :			             Complies / Does not comply

Done by ;			             Checked by	;			                  Authorized by	:

Date:				             Date		 :	                  		      Date                       :

APPENDIX  VI                                     ZYX/CQA/SOP-039/FR-06

TOC SUMMARY SHEET                                    INSCAL----------------------

Instrument ID       : -------------------------------

Location	: -------------------------------

| TEST | OBSERVATION | ACCEPTANCE CRITERIA | STATUS |
| --- | --- | --- | --- |
| By Potassium hydrogen phthalate or By Sucrose |  | Linearity : Correlation coefficient should not be less than 0.98 Standard Deviation : NMT 1.0 RSD : NMT 5.0% | Complies/ Does not Comply |
| By 1,4 Benzoquinone |  | The system is said to be passing the test if the % recovery in between 85% to 115% | Complies/ Does not Comply |

Remarks: All the parameters for the calibration as per predefined acceptance criteria are under limit hence the instrument is suitable

for routine analysis

Remarks :			             Complies / Does not comply

Done by ;			             Checked by	;			                  Authorized by	:

Date:				             Date		 :	                  		      Date                     :