


APPENDIX I	XYZ/CQA/SOP-064/FR-01
                                                                           LOG BOOK OF BOD INCUBATOR
 
	Date of Incubation

	Name of Sample
	Batch / A.R.NO
	Due Date of 
Withdrawal
	Withdrawn on
	Withdrawn by

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	

	
	
	
	
	

	

	
	
	
	
	

	

	
	
	
	
	

	

	
	
	
	
	

	

	
	
	
	
	

	

	
	
	
	
	

	

	
	
	
	
	



                                                                 APPENDIX II                          XYZ/CQA/SOP-064/FR-02                          

                            DAILY TEMPERATURE MONITORING RECORD OF BOD INCUBATOR

BOD Incubator ID:-------------------------                  Set Value : 25°C     
Location:	      :-------------------------                  Limit         : ± 2°C (23 to 27°C)     
	   Date
	Scheduled       Time
	Observed 
 Time
	Temperature
	Observed by
	Checked by
	Remarks
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   XYZ/CQA/SOP-064/FR-03
                                                                    APPENDIX III                            INSCAL…………………                                                   
                                           CALIBRATION DATA SHEET            Page No. 1 of 1
1.0	Instrument Id.:			……………………………………...
2.0	Location:				……………………………………...
3.0	Frequency:				Quarterly
4.0	Date of Calibration:			………………………………………
5.0	Next due date of Calibration:	………………………………………
6.0	PRT Thermometer Id ( Used for Calibration)…………………………..
7.0	Validity of Calibration		………………………………………
8.0      Set Value : 25°C  
9.0	Limit	   : ± 2°C (23 to 27°C)     
	Sr. No.
	Time Interval
	  PRT    Reference Thermometer
	Display on BOD incubator
	Difference between
PRT thermometer & Display on BOD
	Acceptance 
Criteria

	01
	30 min
	
	
	
	PRT & from Display
NMT 1°C

	02
	1 hour
	
	
	
	

	03.
	2 hour
	
	
	
	

	04
	4 hour
	
	
	
	



Remarks:    The above parameter used for the calibration as per predefined acceptance criteria  

                    complies / Does not comply  hence the instrument is suitable / Not suitable  for 
                      
                    routine analysis
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