APPENDIX I

LOG BOOK FOR VISIBLE SPECTROPHOTOMETER

Instrument Id: -----------------------------

Location       : ------------------------------

| Sr. No | Date | Start Time | Name of Sample | Batch No / A. R. No | End Time | Done by | Checked by | Remarks |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

APPENDIX II                                 XYZ/CQA/SOP-089/FR-02

CALIBRATION DATA SHEET	INSCAL-------------------------

Page 1 of 4

1.0	Instrument Id			: ------------------------------------

2.0        Location 			: ------------------------------------

3.0        Frequency of calibration	: HALF YEARLY

4.0	Date of calibration 		: ------------------------------------

5.0	Next due date of calibration: ------------------------------------

6.0	PRELIMINARY CHECK

6.1   	Cleanliness of the Instrument:

Observation: Clean / Not Clean

Corrective action in the case of non-compliance:

6.2	Cleanliness of the cuvettes :

Observation: Ok / Not Ok

Corrective action in the case of non-compliance:

6.3      Humidity and temperature requirement:

Observation: Ok / Not Ok

Corrective action in the case of non-compliance:

6.4       Condition of silica Gel:

Corrective action in the case of non-compliance:

6.5       Instrument is connected to a stabilized power   supply

Observation: Ok / Not Ok

Corrective action in the case of non-compliance:

6.6	Any other maintenance:

6.7	WAVELENGTH ACCURACY	Page 2 of 4

Prepare 1.5M CoCl2.6H2O ----------- g of CoCl2.6H2O (Make-----------Batch / Lot No----------------)    in ------ml of milli Q water.

0.15M CoCl2.6H2O Pipette out ------ml of stock solution to the -------ml volumetric flask and make up the volume with water up to the mark

| Sr, no. | Wavelength in nm | Transmittance |
| --- | --- | --- |
| 1.0 | 600 |  |
| 2.0 | 650 |  |
| 3.0 | 700 |  |
| 4.0 | 750 |  |
| 5.0 | 800 |  |
| 6.0 | 850 |  |

Acceptance Criteria:

Maximum transmittance should be observed at wavelength 700nm±50nm.

Remarks		:	Complies / Does not comply

Done by		:	Checked by	:			Approved by	:

Date			:	Date		:			Date		:

6.1.2    LINEARITY OF ABSORBANCE:	Page 3 of 4

Solution of 1.5M CoCl2.6H2O -------------------------------------------------------------------------

Solution of 0.15M CoCl2.6H2O-------------------------------------------------------------------------

Solution of 0.12M CoCl2.6H2O-------------------------------------------------------------------------

Solution of 0.09M CoCl2.6H2O-------------------------------------------------------------------------

Solution of 0.06M CoCl2.6H2O-------------------------------------------------------------------------

Solution of 0.03M CoCl2.6H2O-------------------------------------------------------------------------

| Sr. no. | concentration | Transmittance at 550nm |
| --- | --- | --- |
| 1 | 0.1500 |  |
| 2 | 0.1200 |  |
| 3 | 0.0900 |  |
| 4 | 0.0600 |  |
| 5 | 0.0300 |  |

Acceptance criteria

Transmittance should be ±2%.

The linearity coefficient ‘r’ obtained from the linearity graph of CoCl2.6H2O for different levels should not be less than 0.99.

Remarks	:                                    Complies / Does not comply.

Done by 	:                                    Checked by	:                                 Approved by:

Date		:                                    Date		:                                 Date		:

12.0     SPECIFICATION OF CELLS:		Page 4 of 4

| Wavelength (nm) | % Transmittance | Acceptance Criteria |
| --- | --- | --- |
|  | Cell |  |
| 400 |  | Not less than 73% |
| 450 |  | Not less than 85% |
| 500 |  | Not less than 87% |
| 550 |  | Not less than 90% |

Remarks	:                                    Complies / Does not comply.

Done by 	:                                    Checked by	:                                 Approved by:

Date		:                                    Date		:                                 Date		:

APPENDIX III                                                 	        XYZ/CQA/SOP-089/FR-03

UV-VIS CALIBRATION SUMMARY SHEEET                                          INSCAL-----------------------------

Page 1 of 2

Instrument ID       : -------------------------------

Location	: -------------------------------

| TEST | ACCEPTANCE CRITERIA |  |  |  | STATUS |  |
| --- | --- | --- | --- | --- | --- | --- |
| Wavelength Accuracy | Wavelength in nm | Observation | Maximum Transmittance | Complies/ Does not Comply |  |
|  | 600 |  |  |  |  |
|  | 650 |  |  |  |  |
|  | 700 |  |  |  |  |
|  | 750 |  |  |  |  |
|  | 800 |  |  |  |  |
|  | 850 |  |  |  |  |
| Linearity of Transmittance | Concentration | Transmittance at 550nm | Acceptance criteria | Complies/ Does not Comply |  |
|  | 0.15M |  | Transmittance should be ±2%. The linearity coefficient ‘r’ obtained from the linearity graph of CoCl2.6H2O for different levels should not be less than 0.99. |  |  |
|  | 0.12M |  |  |  |  |
|  | 0.090M |  |  |  |  |
|  | 0.060M |  |  |  |  |
|  | 0.030M |  |  |  |  |

Page 2 of 2

| Specification of cell | Transmittance | Wavelength in nm | % Transmittance | Complies/ Does not Comply |
| --- | --- | --- | --- | --- |
|  |  | 400 | Not less than 73% |  |
|  |  | 450 | Not less than 85% |  |
|  |  | 500 | Not less than 87% |  |
|  |  | 550 | Not less than 90% |  |

Remarks:	The above parameters used for the calibration as per predefined acceptance criteria complies / Does not comply hence the instrument is suitable / Not suitable for routine analysis.

Done by :                                                Checked BY	:                                                   Approved By    :

Date       : 	                                         Date                     :                                                    Date                  :