APPENDIX                                                                        XYZ/CQA/SOP-053/FR-01

LOG BOOK OF ICP-MS

Instrument Id. No.:	…………………………………….

Location:		……………………………………..

| Sr. No. | Date | Product Name | Batch No/ A. R. No. | Done by | Checked By | Remarks |
| --- | --- | --- | --- | --- | --- | --- |

APPENDIX II                          XYZ/CQA/SOP-053/FR-02

CALIBRATION DATA SHEET	INSCAL ______________

1.0	Instrument Id.:			……………………………………..

2.0	Location:				……………………………………..

3.0       Frequency of calibration:                QUARTERLY

4.0	Date of Calibration:			…………………............................

5.0	Next due date of Calibration:	……………………………………..

6.0       Linearity of Different Concentrations

| S.No. | Heavy Metal | Concentration | Concentration Mean Observed | Intensity Observed | Linearity | Acceptance Criteria |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Arsenic (As) | 2 ppb |  |  |  | 0.99 |
|  |  | 5 ppb |  |  |  |  |
|  |  | 10 ppb |  |  |  |  |
|  |  | 20 ppb |  |  |  |  |
| 2 | Lead (Pb) | 2 ppb |  |  |  | 0.99 |
|  |  | 5 ppb |  |  |  |  |
|  |  | 10 ppb |  |  |  |  |
|  |  | 20 ppb |  |  |  |  |
| 3 | Mercury (Hg) | 2 ppb |  |  |  | 0.99 |
|  |  | 5 ppb |  |  |  |  |
|  |  | 10 ppb |  |  |  |  |
|  |  | 20 ppb |  |  |  |  |
| 4 | Cadmium(Cd) | 2 ppb |  |  |  | 0.99 |
|  |  | 5 ppb |  |  |  |  |
|  |  | 10 ppb |  |  |  |  |
|  |  | 20 ppb |  |  |  |  |

7.0       Performance Check

| S.No. | Analyte | Measured Intensity Mean | Limit of Intensity (cps/10 ppb) |
| --- | --- | --- | --- |
| 1 | Magnesium (Mg) |  | > 40,000 cps/10 ppb |
| 2 | Indium (In) |  | >250,000 cps/10 ppb |
| 3 | Uranium (U) |  | >200,000 cps/10 ppb |

| S.No. | Analyte | Measured Intensity Mean | Limit of Intensity (cps/10 ppb) |
| --- | --- | --- | --- |
| 4 | Cerium Oxide (CeO)/Cerium (Ce) |  | <3% |
| 5 | Barium Ion (Ba++)/ Barium (Ba) |  | <3% |
| 6 | Background 220 |  | <2 cps |

Remarks:				       Complies/Does not Complies

Done By:                                                     Checked By:                                     Approved By:

Date:                                                            Date:                                                Date:

APPENDIX III                                                            XYZ/CQA/SOP-053/FR-03

SUMMARY SHEET                                                    INSCAL________________

Location:		……………………………………..

Instrument Id.:	……………………………………..

LINEARITY AT DIFFERENCE CONCENTRATIONS:

| Concentration | Observations | Acceptance Criteria | Status |
| --- | --- | --- | --- |
| 2 ppb |  | 0.99 | Meet / Does not meet |
| 5 ppb |  | 0.99 | Meet / Does not meet |
| 10 ppb |  | 0.99 | Meet / Does not meet |
| 20 ppb |  | 0.99 | Meet / Does not meet |

STANDARDIZATION:

| Analyte | Measured Intensity Mean | Limit of Intensity (cps/10 ppb) | Status |
| --- | --- | --- | --- |
| Magnesium (Mg) |  | > 40,000 cps/10 ppb | Meet / Does not meet |
| Indium (In) |  | >250,000 cps/10 ppb | Meet / Does not meet |
| Uranium (U) |  | >200,000 cps/10 ppb | Meet / Does not meet |

| Analyte | Measured Intensity Mean | Limit of Intensity (cps/10 ppb) | Status |
| --- | --- | --- | --- |
| Cerium Oxide (CeO)/Cerium (Ce) |  | <3% | Meet / Does not meet |
| Barium Ion (Ba++)/ Barium (Ba) |  | <3% | Meet / Does not meet |
| Background 220 |  | <2 cps | Meet / Does not meet |

Remarks:	All the parameters for the calibration as per predefined acceptance criteria are under limit hence the instrument is suitable for routine analysis

Done By:                                                                                             Checked By:                                                                Approved By:

Date:	                                                                                                 Date:	                                                                           Date: