APPENDIX I                                                               XYZ /CQA/ SOP- 085/FR-01

LOG BOOK OF LC/MS/MS

Instrument ID No.:	------------------------------

Location               :	------------------------------

| Sr.No. | Date | Experiment /Parameters | Sample Details | Batch No. / A. R. No. | Done by | Checked by | Remarks |
| --- | --- | --- | --- | --- | --- | --- | --- |

APPE N D I X I I                                          XYZ / CQA / SOP - 0 85/ FR-02	                                     XYZ  / CQA / SOP-071 /  FR-01

CALIBRATION DATA SHEET          INSCAL----------------------

Page 1 of 11

1.0	Instrument Id:			_________________________

2.0	Location:				_________________________

3.0	Calibration Frequency:		Half Yearly

4.0	Date of Calibration:			_________________________

5.0	Next due date of Calibration:	_________________________

6.0	Calibration of Mass spectrometer

Attach the QQQ

Check tune Report

Remarks:	Complies / Does not comply.

Done by:				Checked by:				Approved by:

Date:					Date:					Date:

7.0	CALIBRATION OF HPLC                                                                           Page 2 of 11

7.1	Pump Calibration

7.1.1	Flow rate Accuracy

7.1.1.1	Calibration of pump “A”

Stop Watch ID No:		---------------------------------

Validity of Calibration:		----------------------------------

Observed Flow Rate             :         Volume Collected / Observed Time

| Set flow | Volume collected (A) | Observed Time(min) (B) |  |  |  |  | Observed Flow Rate (A/B) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Rate ml/min | 10.0 ml | I | II | III | IV | V | I | II | III | IV | V |
| 1.0 |  |  |  |  |  |  |  |  |  |  |  |
| 2.0 |  |  |  |  |  |  |  |  |  |  |  |

Flow Rate Accuracy

| Sr. No. | Flow rate set (ml./min) | Observed flow rate in ml/min. |  |  |  |  |  |  |  | Flow accuracy (%) | Flow precision (% RSD) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | I |  | II |  | III | IV |  | V |  |  |  |
| 1 | 1.000 |  |  |  |  |  |  |  |  |  |  |
| 2 | 2.000 |  |  |  |  |  |  |  |  |  |  |

Acceptance criteria:

Flow Accuracy: 	95 to 105%

Flow Precision:	NMT 2%RSD

Remarks:	Complies / Does not comply.

Done by:				Checked by:				Approved by:

Date:					Date:					Date:

7.1.1.2	Calibration of pump “B”                                                                            Page No. 3 of 11

Stop Watch ID No:		______________________

Validity of Calibration:		______________________

Observed Flow Rate             :         Volume Collected / Observed Time

| Set flow Rate ml/min | Volume collected (A) | Observed Time(min) (B) |  |  |  |  | Observed Flow Rate (A/B) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1.0 | 10.0 ml | I | II | III | IV | V | I | II | III | IV | V |
| 2.0 |  |  |  |  |  |  |  |  |  |  |  |

Flow Rate Accuracy

| Sr. No. |  | Flow rate set (ml./min) |  | Observed flow rate in ml/min. |  |  |  |  |  |  |  |  | Flow accuracy (%) |  | Flow Precision (%RSD) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | I |  | II |  | III |  | IV | V |  |  |  |  |  |
| 1 |  | 1.000 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 |  | 2.000 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

Acceptance criteria:

Flow Accuracy: 95 to 105%

Flow Precision:  NMT 2%RSD

Remarks:	Complies / Does not comply.

Done by:				Checked by:				Approved by:                                                                                                                                               Date:					Date:					Date:

7.1.2	Flow rate Consistency                                                                               Page No. 4 of 11

| Injection Volume | Retention Time |  |  |  |  | %RSD | Acceptance Criteria | Remarks |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Inj.1 | Inj.2 | Inj.3 | Inj.4 | Inj5 |  |  |  |
| 20µL |  |  |  |  |  |  | Not less than 1% |  |

Remarks:	Complies / Does not comply.

Done by:				Checked by:				Approved by:

Date:					Date:					Date:

7.1.3	GRADIENT COMPOSITION ACCURACY:                                                 Page No. 5 of 11

7.1.3.1	Set the following Chromatographic conditions:

Column:			__________________

Solvent for Pump A:		__________________

Solvent for Pump B:		__________________

Flow rate:			__________________

Wavelength:			__________________

Injection volume:		__________________

Run Time:			__________________

Gradient Composition

| Sr. No | Set Concentration ( X % ) | Actual Concentration (X %) | Limit (%) |
| --- | --- | --- | --- |
| 1 | 10 |  | 9 -11 |
| 2 | 50 |  | 49 - 51 |
| 3 | 90 |  | 89 - 91 |

Actual conc. of X% =   absorbance level of set conc.(X%) – absorbance level of set conc. (0%)

absorbance level of set conc. (100%) - absorbance level of set conc. (0%)

_

Actual conc. of 10% = 	--------------------------------------------------------------------------------- =

_

_

Actual conc. of 50% =	---------------------------------------------------------------------------------- =

_

_

Actual conc.  of 90% =	---------------------------------------------------------------------------------- =

_

Acceptance criteria:

At each absorbance level ±1% of set concentration

Remarks:	Complies / Does not comply

Done By:				Checked by:				Approved by:

Date:					Date:					Date:

7.2	AUTOSAMPLER CALIBRATION                                                          Page No. 6 of 11

Balance ID:			____________________

Validity of Calibration:		____________________

7.2.1	Injector Accuracy

7.2.1.2	Chromatographic conditions

Mobile phase:					__________________

Injection volume:				50µL

Weight of Vial before injections W1:		__________________

Weight of vial after 6 injections W2:		__________________

Injection volume                        W3:		    (W1-W2) X 1000

6

(                       )  X 1000

:	            ------------------------------------

6

Acceptance criteria:

50 ± 1.0µL (49-51)

Remarks:	Complies / Does not Comply

Done by:				Checked by:				Approved by:

Date:					Date:					Date:

7.2.2	Injector Precision and Linearity                                                             Page No. 7 of 11

7.2.2.1	Chromatographic condition:

Mobile phase:			___________________

Wavelength:			___________________

Column:			___________________

Injection Volume:		___________________

Run time:			___________________

Flow rate:			___________________

| Injection Volume | %RSD (Area) of replicate injections | Acceptance criteria | Remark |
| --- | --- | --- | --- |
| 5µL |  | NMT1% |  |
| 10 µL |  | NMT1% |  |
| 20 µL |  | NMT1% |  |
| 50 µL |  | NMT1% |  |
| 100 µL |  | NMT1% |  |

Correlation coefficient R2_________________________

Acceptance criteria:

Not less than 0.999

Remarks:	Complies / Does not comply

Done by:				Checked by:				Approved by:

Date:					Date:					Date:

7.2.2.2	CARRYOVER TEST:                                                                                  Page No. 8 of 11

Balance ID:			_____________________

Validity for Calibration:	_____________________

7.2.2.2.1 Chromatographic conditions

Mobile phase:			_____________________

Column type:				_____________________

Flow rate:			________________ml/min

Detector wavelength:		________________nm

Column oven temperature:	________________°C

Injection volume:		_________________µl

7.2.2.2.2	Preparation of standard solution:

Wt. of Caffeine (Lot No/Batch No. ________________)

___________ g dissolved and diluted to __________ ml with mobile phase.

| Injection Volume | Area |
| --- | --- |
| Pre blank |  |
| Caffeine Solution |  |
| Post blank |  |

%Carry over =	 Area count of post blank  - Area count of pre blank

Area count of caffeine solution

_

=	-----------------------------------------------------------------------

Acceptance Criteria:

NMT 0.05%

Remarks:	Complies / Does not comply.

Done by:				Checked by:				Approved by:

Date:						Date:					Date:

7.3	DETECTOR CALIBRATION                                                                     Page No. 9 of 11

7.3.1	LINEARITY OF DETECTOR RESPONSE

Balance ID:			____________________

Validity for Calibration:	____________________

7.3.1.2 Chromatographic conditions

Mobile phase:			____________________

Column type:			____________________

Flow rate:			_______________ml/min

Detector wavelength:		_______________nm

Column oven temperature:	_______________°C

Injection volume:		100ppm, 50ppm, 25ppm, 10ppm, 5ppm.

Preparation of standard solution:

Wt. of Caffeine (Lot No/Batch No.-------------------------------)

-----------------g dissolved and diluted to ----------------- ml with mobile phase. (1000ppm)

Dilute _______	ml of solution (1000ppm) to	________	ml with mobile phase (100ppm)

Dilute _______	ml of solution (100ppm) to	________	ml with mobile phase (500ppm)

Dilute _______	ml of solution (50ppm) to	________	ml with mobile phase (25ppm)

Dilute _______	ml of solution (25ppm) to	________	ml with mobile phase (10ppm) Dilute _______	ml of solution (10ppm) to	________	ml with mobile phase (5ppm)

Acceptance criteria:

The Linearity coefficient ‘r’ obtained from linearity graph of caffeine should not be less than 0.999.

Remarks:	Complies / Does not comply.                                                                                                 Done by:				Checked by:				Approved by:

Date:					Date:					Date:

7.3.2	WAVELENGTH ACCURACY                                                                   Page No. 10 of 11

7.3.2.1	Using Caffeine

Lot No/Batch No of Caffeine:		__________________

Mobile Phase:				__________________

Column:				__________________

Flow rate:				__________________

Caffeine in water:			__________________

Injection Volume:			__________________

Observation:

| Sr.No | Wavelength (in nm) | Observed Wavelength(in nm) |
| --- | --- | --- |
| 01 | 271 |  |
| 02 | 272 |  |
| 03 | 273 |  |
| 04 | 274 |  |
| 05 | 275 |  |

Acceptance Criteria:

273 ±1nm

Remarks:	Complies / Does not comply.

Done by:				Checked by:				Approved by:

Date:					Date:					Date:

7.4	COLUMN OVEN TEMPERATURE                                                            Page No. 11 of 11

Thermometer ID:		__________________

Validity of Calibration:		__________________

| Set Temp | Observed temp | Limit (± 2°C) | Remark |
| --- | --- | --- | --- |
| 40°C |  | 38 - 42°C |  |
| 50°C |  | 48 - 52°C |  |

Acceptance Criteria:

The observed reading and standard thermometer reading do not differ by ±2°C from the set temperature.

Remarks:	Does not comply

Done by:				Checked by:				Approved by:

Date:					Date:					Date:

APPENDIX III                                                                 XYZ /CQA/SOP-085/FR-03

LC/MS/MS CALIBRATION SUMMARY SHEET                                            Page No. 1 of 3

Instrument ID:  	---------------------------

Location:		----------------------------

| Test Name | Observation |  |  |  |  |  | Acceptance Criteria | Status |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CALIBRATION OF HPLC |  |  |  |  |  |  |  |  |
| Calibration of the pump flow rate | Flow Rate (ml/min) | Flow Accuracy |  |  | Flow Precision |  | Flow accuracy of the set flow rate (95.0 – 105%) Flow precision RSD NMT 2.0% | Complies/ Does not Comply |
|  |  | A |  | B | A | B |  |  |
|  | 1.0 |  |  |  |  |  |  |  |
|  | 2.0 |  |  |  |  |  |  |  |
| Flow rate Consistency |  |  |  |  |  |  | RSD of RT NMT1% | Complies/ Does not Comply |
| Gradient Composition Accuracy | 10 % |  | 50% |  | 90% |  | At each absorbance level ± 1% of set concentration | Complies/ Does not Comply |

Page No. 2 of 3

| Test Name | Observation | Acceptance Criteria | Status | Test Name |
| --- | --- | --- | --- | --- |
| Column Oven Calibration | 40°C | 50°C | ± 2°C of set value | Complies/ Does not Comply |
| Injector Precision |  |  | RSD of Area NMT1% | . Complies/ Does not Comply |
| Injector Linearity | Correlation coefficient `r’ = --------------- |  | NLT 0.999 | Complies/ Does not Comply |
| Injector Accuracy |  |  | 50± 1.0 µL(49 µL to51 µL) | Complies/ Does not Comply |
| Wavelength accuracy | Wavelength = ------------------ nm (Using Caffeine) |  | NMT 273.0 nm ± 1.nm | Complies/ Does not Comply |
| Linearity of Detector Response | Correlation coefficient `r’ = --------------- |  | NLT 0.99 | Complies/ Does not Comply |

Page No. 3 of 3

| Test Name | Observation | Acceptance Criteria | Status |
| --- | --- | --- | --- |
| Carry over Test | Carryover = ------------------- % | Carry over of Caffeine NMT 0.05% of area of 100 µl injection of 250ppm Standard Solution. | Complies/ Does not Comply |

Remarks:	All the parameters for the calibration as per predefined acceptance criteria complies / Does not comply hence the instrument is suitable / Not Suitable for routine analysis.

Done by:								Checked by:						Approved by:

Date:								Date:							Date:

APPENDIX IV                       XYZ /CQA/SOP-085/FR-04

DAILY CHECKPOINTS

| ##### Date | Cleaning of Source | Column Cleaning | Vacuum Pressure (3.0 e-5 torr to 4.0e-5 torr) |  | Gases Pressures |  |  |  |  | ###### Done By | ###### Checked By |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | Nitrogen (2 kg/cm2) |  |  | Nitrogen Generator (80-100 psi) |  |  |  |  |

APPENDIX V	XYZ /CQA/SOP-085/FR-05

WEEKLY CHECKPOINTS

| Date | Oil Level of Vacuum Pump | Done By | Checked By |
| --- | --- | --- | --- |

APPENDIX VI	XYZ /CQA/SOP-085/FR-06

| Date | Calibration of Mass Spectrometer | Calibration of Pump | Calibration of Oven | Calibration of UV - Vis Detector | Calibration of Auto sampler | Done By | Checked By |
| --- | --- | --- | --- | --- | --- | --- | --- |

HALF YEARLY CHECKPOINTS