## OAE
Auditory Function Screening Device – Neonatal Otoacoustic Emission (OAE) Test System
Bidder shall complete this form to provide responses about their proposed medical equipment. Each response by Bidder shall be supported with reference document,e.g: Catalogue, Equipment Technical Data Sheet, etc., and each reference shall be highlighted and/or labelled.
UMDNS: 20167 Auditory Function Screening Devices, Newborn (Otoacoustic) Total Quantity: 32 units |  | Equipment Location (Region): Quantity SIHREN Indonesia (Nationwide): 32 units | Description: Auditory function screening devices are designed to detect hearing loss in neonates and infants through the measurement of otoacoustic emissions (OAE). These devices assess the function of the cochlear outer hair cells by detecting sound energy (emissions) generated within the cochlea in response to an external auditory stimulus. OAE screening systems typically consist of an audio frequency stimulus generator, ear probe with miniature loudspeaker and microphone, connecting cables, and a computerized processing unit with integrated display and/or printer for data visualization and report generation.
Row no. | PROPOSED DEVICE INFORMATION | REQUIREMENTS | BIDDER RESPONSE
|  |  | Device Details [Please specify how the proposed device meets each requirement] |  |  | Source Data [Please provide the name and page number of the relevant supporting document(s)] | Source Data URL [Please provide the URL of the source document(s), if available]
1 | Brand/Manufacturer Name | Bidder to specify | Enter Responses Here! |  |  | Enter Responses Here! | Enter Responses Here!
2 | Model Name | Bidder to specify | Enter Responses Here! |  |  | Enter Responses Here! | Enter Responses Here!
3 | Country of Origin | Bidder to specify | Enter Responses Here! |  |  | Enter Responses Here! | Enter Responses Here!
4 | Year of 1st Production for Model Proposed | Bidder to specify | Enter Responses Here! |  |  | Enter Responses Here! | Enter Responses Here!
5 | Brochure/Catalogue | Bidder shall provide product brochure in English. Both Soft copy and Hard copy are required. | Enter Responses Here! |  |  | Enter Responses Here! | Enter Responses Here!
6 | Regulatory Approval | Required; Bidder to specify whether the offered device has an Indonesian Product Approval License (Nomor Izin Edar, NIE). The NIE is not required at the bidding stage but must be obtained prior to contract signing. If the device does not yet have an NIE, the bidder must submit valid evidence of compliance with an internationally recognized regulatory standard, such as U.S. FDA clearance, CE Mark certification (MDR EU 2017/745), or an equivalent health, safety, and environmental standard approval. | Enter Responses Here! |  |  | Enter Responses Here! | Enter Responses Here!
7 | Picture of Model Proposed | Bidder to attach a picture of model proposed | Enter Responses Here! |  |  | Enter Responses Here! | Enter Responses Here!
8 | CLINICAL APPLICATIONS/PROCEDURES | REQUIREMENTS | BIDDER RESPONSE
9 |  |  | Yes [√] | No [x] | Device Details [Please specify how the proposed device meets each requirement] | Source Data [Please provide the name and page number of the relevant supporting document(s)] | Source Data URL [Please provide the URL of the source document(s), if available]
10 | To evaluate cochlear (outer hair cell) function through otoacoustic emissions (OAE) testing. | Required; For newborn and infant hearing screening. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
11 | TECHNICAL SPECIFICATIONS/FEATURES | REQUIREMENTS | BIDDER RESPONSE
12 |  |  | Yes [√] | No [x] | Device Details [Please specify how the proposed device meets each requirement] | Source Data [Please provide the name and page number of the relevant supporting document(s)] | Source Data URL [Please provide the URL of the source document(s), if available]
13 | Note: Unless otherwise stated, all specifications constitute minimum performance requirements. Devices that meet or exceed these requirements are acceptable. Bidders must provide documentary technical evidence (e.g., user manual, service manual, technical brochure, signed declaration from the manufacturer, or video evidence), with clear annotations to the source data, to support all claims of compliance.
14 | General
15 | Configuration | Required; Portable, handheld. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
16 | Operating System / Automation | Required; Automated screening system with embedded software that guides users through test setup, execution, and result display to enable reliable screening with minimal operator training. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
17 | General Screening and Operational Features
18 | Screening Capabilities | Required; Device must support otoacoustic emission (OAE)–based hearing screening for neonates and infants. Bidder to ensure the following two testing modes are supported: a) Transient Evoked OAE (TEOAE) b) Distortion Product OAE (DPOAE) |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
19 | Screening Capabilities | Bidder to specify any other screening modalities offered (if any). |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
20 | Operational Modes | Required; Alternating. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
21 | Calibration | Required; Automatic or self-calibration prior to each test. Bidder to specify. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
22 | Transient Evoked Otoacoustic Emissions (TEOAE)
23 | Stimulus Type | Required; Click or tone-burst stimulus. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
24 | Stimulus Intensity | Required; 80 decibels peak equivalent sound pressure level (dB peSPL). |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
25 | Frequency Range | Required; 1.5 – 4 kilohertz (kHz). |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
26 | Test Time | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
27 | Referral Rate | Required; ≤ 4%. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
28 | Distortion Product Otoacoustic Emissions (DPOAE)
29 | Stimulus Type | Required; Dual-tone (f₁ and f₂) stimulus. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
30 | Stimulus Intensity | Required; 40 – 70 decibels sound pressure level (dB SPL). Bidder to specify exact range. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
31 | Frequency Range | Required; 1.5 – 12 kHz. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
32 | Test Time | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
33 | Referral Rate | Required; ≤ 4%. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
34 | Test Outcome and Display Features
35 | Result Presentation | Required; Pass/Refer indication displayed on screen at the end of each test. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
36 | Display Information | Bidder to specify full list of displayed information (e.g., signal level, test progress, ear tested, test quality, etc.). |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
37 | User Interface | Required; Bidder to specify user interface type and describe key operational features (e.g., menu navigation, on-screen prompts, icons). |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
38 | Data Management and Connectivity
39 | Device Data Storage Capacity | Required; ≥ 500 test results. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
40 | Data Transfer Capabilities | Required; Capability for data transfer via USB and/or wireless (Bluetooth or Wi-Fi). Bidder to specify all available options. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
41 | Report Types | Bidder to specify available formats (e.g., waveform, tabular, text, etc.). |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
42 | Alarm and Safety Functions
43 | Alarm Indication | Required; Audible and/or visual alarms to alert users to test or device conditions affecting screening accuracy. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
44 | Probe Fit Error Alarm | Required |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
45 | High Noise or Artifact Detection Alarm | Bidder to specify if offered. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
46 | Device Malfunction Alarm | Bidder to specify if offered. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
47 | Test Pause Function | Required; Automatic and manual test pause available. Bidder to specify all conditions that trigger automatic test pause (e.g., high ambient, poor probe fit, etc.). |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
48 | Cleaning and Disinfection
49 | Type of Disinfectant Used for Infection Control | Required; Free of corrosive ingredients |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
50 | Compatible Agents | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
51 | Cleaning Instruction/Guideline for Clinicians/Users | Bidder to provide |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
52 | ACCESSORIES REQUIRED FOR EACH CLINICAL APPLICATION/PROCEDURES | REQUIREMENTS | BIDDER RESPONSE
53 |  |  | Yes [√] | No [x] | Device Details [Please specify how the proposed device meets each requirement] | Source Data [Please provide the name and page number of the relevant supporting document(s)] | Source Data URL [Please provide the URL of the source document(s), if available]
54 | Carrying Case | Required; One (1) per unit. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
55 | OAE Probe | Required; One (1) complete probe per unit, including probe tip and connecting cable. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
56 | Other Proprietary Accessories Required for Full Functionality of the Device (if any) | Bidders must specify and supply any additional proprietary accessories not listed above that are required to ensure the full functionality and intended performance of the proposed device. These accessories are to be included in the bid. If applicable, bidders must specify the quantity and cost (only in the financial part) of each accessory in the “Proposed Accessory” sheet. If no additional proprietary accessories are required, bidders must indicate “N/A.” |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
57 | START-UP CONSUMABLES, if any | REQUIREMENTS | BIDDER RESPONSE
58 |  |  | Yes [√] | No [x] | Device Details [Please specify how the proposed device meets each requirement] | Source Data [Please provide the name and page number of the relevant supporting document(s)] | Source Data URL [Please provide the URL of the source document(s), if available]
59 | Disposable Ear Tips, neonatal size | Required; Bidder to supply a starter kit in a quantity required for 1,000 typical procedures. Bidder to specify unit quantity in the "Proposed Consumables" sheet. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
60 | Other Proprietary Consumables Required for Full Functionality of the Device (if any) | Bidders must specify and supply any additional proprietary consumables not listed above that are required to ensure the full functionality and intended performance of the proposed device. These consumables must be included in the starter kit, and a fixed price must be quoted for a five-year supply upon full commissioning of the device. If applicable, bidders must specify the quantity and cost (only in the financial part) of each consumable required for 1,000 typical procedures in the “Proposed Consumables” sheet. If no additional proprietary consumables are required, bidders must indicate “N/A.” |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
61 | Other Non-Proprietary Consumables Required for Full Functionality of the Device (if any) | Bidders must specify any additional non-proprietary consumables not listed above that are required to ensure the full functionality and intended performance of the proposed device. This information is for reference purposes only (i.e., these consumables will not be included in the bid or supplied upon commissioning of the device). If no additional non-proprietary consumables are required, bidders must indicate “N/A.” |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
62 | DEVICE PHYSICAL ATTRIBUTES | REQUIREMENTS | BIDDER RESPONSE
63 |  |  | Yes [√] | No [x] | Device Details [Please specify how the proposed device meets each requirement] | Source Data [Please provide the name and page number of the relevant supporting document(s)] | Source Data URL [Please provide the URL of the source document(s), if available]
64 | Unit Dimension, H x W x D, cm | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
65 | Unit Weight, kg | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
66 | MECHANICAL & ELECTRICAL (M&E) REQUIREMENTS | REQUIREMENTS | BIDDER RESPONSE
67 |  |  | Yes [√] | No [x] | Device Details [Please specify how the proposed device meets each requirement] | Source Data [Please provide the name and page number of the relevant supporting document(s)] | Source Data URL [Please provide the URL of the source document(s), if available]
68 | Power Source | Required; Battery operated. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
69 | Battery Type | Required; Rechargable. Bidder to specify if batteries are user-replacable. Bidder requirements: a) If batteries are user-replaceable: bidder to provide two (2) battery sets per unit (one primary, one spare) b) If batteries are non-removable: bidder to ensure device is equipped with one (1) integrated battery. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
70 | Battery Operating Time, hr | Required; ≥ 8 hours |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
71 | Battery Recharging Time, hr | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
72 | Battery Indicator | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
73 | Battery Charger/Adapter and Cables | Required; Bidder to provide one (1) complete set per unit, meeting the following requirements: a) Electrical Requirements: 230 ± 10 V AC, 50Hz, Single-Phase; and b) Power Plug: Hospital-grade Type F plug, compliant with Indonesian regulations. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
74 | IT REQUIREMENTS | REQUIREMENTS | BIDDER RESPONSE
75 |  |  | Yes [√] | No [x] | Device Details [Please specify how the proposed device meets each requirement] | Source Data [Please provide the name and page number of the relevant supporting document(s)] | Source Data URL [Please provide the URL of the source document(s), if available]
76 | Compatible Data Management System (Middleware) | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
77 | Data Port | Required; Bidder to specify available physical interfaces (e.g., USB, micro-USB, RS-232) for data download and service access. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
78 | Communication Protocol | Bidder to specify supported communication standards/data formats for integration with external systems. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
79 | System Upgradeable in Future | Required; Software and hardware updates and upgrades to be provided at no additional cost. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
80 | Cybersecurity Considerations
81 | Installation of Current Release Software Version | Required; Device must be supplied with the latest stable software release. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
82 | Anti-malware / Anti-virus Compatibility | Bidder to specify if the device includes or is compatible with anti-malware or anti-virus protection software. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
83 | Deactivation of Unused Communication Channels | Bidder to specify if the device offers the ability to disable (software and/or hardware) unused or unnecessary communication ports (e.g. USB, LAN, Wi-Fi, Bluetooth) to enhance data security. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
84 | Confidentiality and Security Measures | If applicable, bidder to specify encryption, user authentication, and data protection mechanisms. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
85 | ARTIFICIAL INTELLIGENCE (AI) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER RESPONSE
86 |  |  | Yes [√] | No [x] | Device Details [Please specify how the proposed device meets each requirement] | Source Data [Please provide the name and page number of the relevant supporting document(s)] | Source Data URL [Please provide the URL of the source document(s), if available]
87 | N/A | N/A
88 | TELEMEDICINE (TM) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER RESPONSE
89 |  |  | Yes [√] | No [x] | Device Details [Please specify how the proposed device meets each requirement] | Source Data [Please provide the name and page number of the relevant supporting document(s)] | Source Data URL [Please provide the URL of the source document(s), if available]
90 | Does the equipment have telemedicine capability or can it be used in telemedicine applications? | Bidder to specify telemedicine capabilities |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
91 | Asynchronous telemedicine capabilities. Store-and-forward technology is asynchronous in that it allows for the images or data to be interpreted at a later time and eliminates the need for the patient and physician to be in the same room or venue. (i.e., through store and forward, in which patient information, such as an image or data, "stored" in the originating facility's (i.e., where the patient is) computer system is sent securely ("forwarded") to the telehealth provider at the distant site (i.e., where the provider is) for review.) | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
92 | Remote patient monitoring capabilities. Remote patient monitoring, in which vital sign information (e.g., blood pressure) and other data (e.g., blood glucose levels) are collected by monitoring devices, usually at home, and transferred to the patient's physician. This technology is used often in the care of patients with chronic conditions. | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
93 | Mobile health (mHealth) integration capabilities. Mobile health (mHealth), in which healthcare and public health information are communicated through mobile devices. The information may include general educational information, targeted texts, and notifications about disease outbreaks. | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
94 | Other telemdicine capabilities. | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
95 | Video Capabilities (if available)
96 | Are video capabilities needed? If so, specify if they are included in the solution or offered through 3rd party add-on? | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
97 | Audio capabilities (if available)
98 | Is audio functionality available? | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
99 | MEDICAL GASES REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER RESPONSE
100 |  |  | Yes [√] | No [x] | Device Details [Please specify how the proposed device meets each requirement] | Source Data [Please provide the name and page number of the relevant supporting document(s)] | Source Data URL [Please provide the URL of the source document(s), if available]
101 | N/A | N/A
102 | ROOM REQUIREMENT | REQUIREMENTS | BIDDER RESPONSE
103 |  |  | Yes [√] | No [x] | Device Details [Please specify how the proposed device meets each requirement] | Source Data [Please provide the name and page number of the relevant supporting document(s)] | Source Data URL [Please provide the URL of the source document(s), if available]
104 | Temperature Requirements | The ambient temperature in the area ranges from 22°C to 32°C. The device must be fully operational within this temperature range. Bidder to specify the full operational temperature range of the proposed device. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
105 | Humidity Requirements | The relative humidity in the area ranges from 50% to 80%. The device must be non-condensing and fully operational within this humidity range. Bidder to specify the full operational humidity range of the proposed device. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
106 | Other | Bidder to specify |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
107 | MAINTENANCE & SERVICE REQUIREMENTS | REQUIREMENTS | BIDDER RESPONSE
108 |  |  | Yes [√] | No [x] | Device Details [Please specify how the proposed device meets each requirement] | Source Data [Please provide the name and page number of the relevant supporting document(s)] | Source Data URL [Please provide the URL of the source document(s), if available]
109 | Inspection, Testing & Commissioning | Required; Inclusive of vendor representative attendance, with test equipments and supplies. Bidder must carry out Inspection, Testing, and Commissioning, including Electrical Safety Testing (where applicable), using calibrated test equipment, tools, and a procedural checklist. Bidder shall inspect and test all equipment supplied to the hospital in accordance with the manufacturer’s recommended procedures. Note for Bidders: These services shall be included under Related Services and priced in the Price Schedules – Related Services and Completion Schedule. The Bidder shall ensure that all costs associated with the provision of these services are fully included in the Contract Price. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
110 | User Training | Required; Twice. First training: delivered within 2 weeks upon successful Acceptance Second training: refresh course, delivered approximately three months later. Note for Bidders: These services shall be included under Related Services and priced in the Price Schedules – Related Services and Completion Schedule. The Bidder shall ensure that all costs associated with the provision of these services are fully included in the Contract Price. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
111 | Duration of User Training Session (Time) | Bidder to specify the duration of training session |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
112 | Operator Manual (Instruction For Use) and Service Manual | Required; Bidder to provide in Hardcopy and/or Softcopy upon successful Inspection, Testing & Commissioning. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
113 | Warranty and Comprehensive Operations and Maintenance Contract Duration | Required; To commence upon successful Acceptance Testing & Commissioning. Warranty and Comprehensive Operations and Maintenance Contract Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation; inclusive of all supplied items provided in the offer. Note for Bidders: These services shall be included under Related Services and priced in the Price Schedules – Related Services and Completion Schedule. The Bidder shall ensure that all costs associated with the provision of these services are fully included in the Contract Price. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
114 | Extension of Warranty and Comprehensive Operations and Maintenance Contract Duration | N/A
115 | Manufacturer Recommended Annual Preventive Maintenance Interval | Required; ≥ 2 times per year. Must be inclusive of all supplied items provided in the offer. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
116 | Comprehensive Service Contract Duration, Post Warranty | Required; To commence upon end of Warranty Duration. Comprehensive Service Contract Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation. Inclusive of all supplied items provided in the offer. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
117 | Response Time, Remote Technical Support (Phone/Email) | Required; Within one (1) hour upon notification from user, inclusive of weekends and public holidays. Applicable for both Warranty Duration and Comprehensive Service Contract Duration. Note for Bidders: This service shall be included and priced within the warranty and the comprehensive operation and maintenance contract under the Price and Completion Schedule for Related Services. Bidder to ensure that all costs associated with the provision of this service are fully included in the Contract Price. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
118 | Response Time, On-site Technical Support | Required; Within 48 hours upon notification from user, inclusive of weekends and public holidays. Applicable for both Warranty Duration and Comprehensive Service Contract Duration. Note for Bidders: This service shall be included and priced within the warranty and the comprehensive operation and maintenance contract under the Price and Completion Schedule for Related Services. Bidder to ensure that all costs associated with the provision of this service are fully included in the Contract Price. |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
119 | Maximum Turnaround Time for Corrective Maintenance | Required; ≤ 120 hours (with penalty system). |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
120 | Uptime Guarantee | Required; Bidder to ensure a minimum average of 95 percent uptime, calculated across all fully commissioned equipment over a six-monthly period |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
121 | Assured Availability of Spare Parts and Proprietary Consumables | Required; ≥ 10 years post Acceptance Testing and Commissioning Date |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!
122 | Installation Base (Specific to this model) | Please provide installation base (facility name and location), year of installation |  |  | Enter Responses Here! | Enter Responses Here! | Enter Responses Here!

## Hazards and Recalls
Notification of Equipment Hazards and/or Recalls
Please list down the Hazards and/or Recalls of the equipment that you proposed and the field corrective actions
List (Reported Hazards, Alerts, RECALL): | Please provide clarification of action taken:
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!

## Proposed Accessories & Price
Usulan Aksesoris dan Harga (5 Tahun)
Procurement of Auditory Function Screening Device – OAE Test System
Termasuk deskripsi dan harga untuk semua aksesoris yang ditawarkan oleh Perusahaan Anda dan diperlukan untuk fungsi penuh perangkat
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
Nomor Item Baris [Penyedia menambhakan angka sesuai kebutuhan] | Nama Aksesoris/Deskripsi [Penyedia memasukkan jawaban ditiap baris di bawah] | Merk [Penyedia memasukkan jawaban ditiap baris di bawah] | Tipe [Penyedia memasukkan jawaban ditiap baris di bawah] | No.Katalog [Penyedia memasukkan jawaban ditiap baris di bawah] | Hitungan Satuan/Kuantitas [Kolom ini telah diisi sebelumnya. Mohon untuk tidak merubah, kecuali aksesoris baru yang diusulkan] | Mata uang dari Harga Satuan (Kol. 8) and Jumlah Harga (Kol. 9) [Kolom ini sudah terisi otomatis. Mohon untuk tidak merubah] | Harga Satuan [Penyedia memasukkan jawaban ditiap baris di bawah] | Jumlah Harga [Tolom ini telah diisi sebelumnya. Mohon untuk tidak merubah,
1 | Carrying Case | Enter response here! | Enter response here! | Enter response here! | 1 | IDR | Enter response here! | 0
2 | OAE Probe | Enter response here! | Enter response here! | Enter response here! | 1 | IDR | Enter response here! | 0
3 | [5 empty] | IDR |  | 0
4 | [5 empty] | IDR |  | 0
5 | [5 empty] | IDR |  | 0
6 | [5 empty] | IDR |  | 0
7 | [5 empty] | IDR |  | 0
8 | [5 empty] | IDR |  | 0
9 | [5 empty] | IDR |  | 0
10 | [5 empty] | IDR |  | 0
11 | [5 empty] | IDR |  | 0
12 | [5 empty] | IDR |  | 0
13 | [5 empty] | IDR |  | 0
14 | [5 empty] | IDR |  | 0
15 | [5 empty] | IDR |  | 0
[5 empty] | Harga Total (jumlah dari kolom 9) |  | IDR | 0

## Proposed Consumables & Price
Usulan Bahan Habis Pakai dan Harga (5 Tahun)
Procurement of Auditory Function Screening Device – OAE Test System
Bahan habis pakai yang tercantum dalam tabel berikut ini harus disediakan untuk setiap unit yang dipasang dalam jumlah kit awal yang ditentukan, diikuti dengan harga tetap per set yang dipesan dari jumlah yang ditentukan tersebut untuk periode 5 tahun setelah berhasil uji penerimaan dan komisioning. (CATATAN: Nama merk, jika digunakan, tidak menunjukkan preferensi khusus dari Pembeli dan hanya diberikan sebagai contoh jenis produk yang diperlukan)
Informasi Umum | [6 empty] | Starter kit |  |  | Perlengkapan Tambahan |  |  | Total Perkiraan Biaya untuk 5 Tahun |  |  | Informasi Tambahan
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 |  | 10 | 11 |  | 12 | 13 |  | 14
Nomor Item Baris [Penyedia menambhakan angka sesuai kebutuhan] | Name Bahan Habis Pakai/Deskripsi [Penyedia memasukkan jawaban ditiap baris di bawah] | Merk [Penyedia memasukkan jawaban ditiap baris di bawah] | Tipe [Penyedia memasukkan jawaban ditiap baris di bawah] | No.Katalog [Penyedia memasukkan jawaban ditiap baris di bawah] | Syarat penyimpanan bahan habis pakai yang ditentukan [Peserta lelang diharuskan mengisi jawaban pada setiap baris di bawah ini] | Umur simpan minimum yang wajib untuk barang-barang konsumsi yang ditentukan [Peserta lelang diharuskan mengisi jawaban pada setiap baris di bawah ini] | Jumlah yang akan disediakan sebagai bagian dari paket awal yang akan disediakan untuk setiap unit yang dipasang. Biaya yang sesuai dengan paket awal ini merupakan bagian dari harga penawaran. [Penawar diharuskan mengisi tanggapan pada setiap baris di bawah ini, jika belum diisi sebelumnya] | Mata Uang [Kolom ini telah diisi secara otomatis. Mohon jangan diedit] | Harga starter kit. Harga ini sudah termasuk harga penawaran [Penyedia memasukkan jawaban di tiap baris di bawah] | Kuantitas yang harus diajukan oleh Penyedia (hanya pada bagian keuangan) dengan harga tetap (termasuk pengiriman) untuk periode 5 tahun setelah berhasil uji coba penerimaan dan komisioning. Biaya yang sesuai untuk paket bahan habis pakai ini tidak termasuk dalam harga penawaran/kontrak, tetapi menjadi bagian dari evaluasi keuangan penawaran. [Penyedia memasukkan jawaban di tiap baris di bawah, jika tidak sudah terisi sebelumnya] | Mata Uang [Kolom ini telah diisi secara otomatis. Mohon jangan diedit] | Harga per set tambahan bahan habis pakai (termasuk pengiriman) dengan jumlah unit tercantum di kolom K, berlaku tetap selama 5 tahun setelah uji coba dan serah terima berhasil. Penawaran mengikat ini tidak termasuk dalam harga penawaran/kontrak, tetapi menjadi bagian dari evaluasi keuangan penawaran. [Penawar diharuskan mengisi tanggapan di setiap baris di bawah ini] | Faktor untuk menghitung total biaya proyeksi bahan habis pakai selama periode 5 tahun [Kolom ini akan diisi otomatis. Mohon jangan diedit] | Mata Uang [Kolom ini telah diisi secara otomatis. Mohon jangan diedit] | Total biaya perkiraan bahan habis pakai selama periode 5 tahun (kolom M x kolom N). Biaya ini tidak termasuk dalam harga penawaran/kontrak, tetapi menjadi bagian dari evaluasi keuangan penawaran. [Kolom ini akan diisi otomatis. Mohon jangan diedit] | Penawar diharuskan untuk menyediakan informasi tambahan yang relevan mengenai produk yang ditawarkan (spesifikasi teknis, sertifikat kualitas, dll.) [Penawar diharuskan untuk mengisi tanggapan pada setiap baris di bawah ini, jika diperlukan/berlaku]
Disposable Ear Tips, neonatal size
1 | Disposable Ear Tips, neonatal size | Enter response here! | Enter response here! | Enter response here! | N/A | N/A | Enter response here! [Bidder to specify the unit quantity required for 1,000 typical procedures] | IDR | Enter response here! | Enter response here! [Bidder to specify the unit quantity required for 1,000 typical procedures] | IDR | Enter response here! | 5 | IDR | 0 | Enter response here! [Only if needed/applicable]
Other Proprietary Accessories Required for Full Functionality of the Device (if any)
2 | Enter response here! [Only if needed/applicable] | Enter response here! [Only if needed/applicable] | Enter response here! [Only if needed/applicable] | Enter response here! [Only if needed/applicable] | N/A | N/A | Enter response here! [If provided, bidder to specify the unit quantity of each specified consumable required for 1,000 typical procedures] | IDR | Enter response here! [Only if needed/applicable] | Enter response here! [If provided, bidder to specify the unit quantity of each specified consumable required for 1,000 typical procedures] | IDR | Enter response here! [Only if needed/applicable] | 5 | IDR | 0 | Enter response here! [Only if needed/applicable]
Other Non-Proprietary Consumables Required for Full Functionality of the Device (if any; not included in the starter kit, for information purposes only) | [16 empty] | wce
3 | Enter response here! [Only if needed/applicable] | Enter response here! [Only if needed/applicable] | Enter response here! [Only if needed/applicable] | Enter response here! [Only if needed/applicable] | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Enter response here! [Only if needed/applicable]
Total Harga [Bagian ini akan diisi otomatis. Mohon jangan diedit]
|  |  | Mata Uang [Kolom ini akan terisi secara otomatis. Mohon jangan dirubah] | Harga [Kolom ini akan terisi secara otomatis. Mohon jangan dirubah]
TOTAL HARGA: Starter kit (JUMLAH of Col. 9) |  |  | IDR | 0 | [ Catatan: Harga-harga ini harus sesuai dengan Jadwal Harga dan Jadwal Penyelesaian - Layanan Terkait, Layanan No.5 Bahan Konsumsi Proprietary ] (Starter Kit)]
Total biaya perkiraan bahan habis pakai selama periode 5 tahun (jumlah Kolom 13) |  |  | IDR | 0 | [ Catatan: Harga-harga ini harus sesuai dengan Jadwal Harga dan Jadwal Penyelesaian - Layanan Terkait, Layanan No.5 Bahan Konsumsi Proprietary (Set Tambahan)]

## Proposed Serv.Contract PW
Usulan Kontrak Layanan Pemeliharaan Pasca-Garansi (Tahun 6 ke 10)
Procurement of Auditory Function Screening Device – OAE Test System
Merk [Penyedia memasukkan jawaban ditiap baris di bawah] | Tipe [Penyedia memasukkan jawaban ditiap baris di bawah] | Hitungan Satuan/Kuantitas [Penyedia memasukkan jawaban ditiap baris di bawah] | Frekuensi PPM (kali/tahun) [Penyedia memasukkan jawaban ditiap baris di bawah] | Kontrak Layanan Komprehensif
Harus Mencakup: - Jam Kerja - Biaya Jasa (Tenaga Kerja) serta Biaya Transportasi dan Akomodasi - Suku Cadang
|  |  |  | Rincian Kontrak Layanan [Penyedia memasukkan jawaban ditiap baris di bawah] | Mata Uang Harga Penawaran untuk Tahun ke 6 - 10 [Kolom ini terisi otomatis,mohon untuk tidak merubah] | Total untuk tahun ke 6 [Penyedia memasukkan jawaban di setiap baris di bawah ini] | Total untuk tahun ke 7 [Penyedia memasukkan jawaban di setiap baris di bawah ini] | Total untuk tahun ke 8 [Penyedia memasukkan jawaban di setiap baris di bawah ini] | Total untuk tahun ke 9 [Penyedia memasukkan jawaban di setiap baris di bawah ini] | Total untuk tahun ke 10 [Penyedia memasukkan jawaban di setiap baris di bawah ini]
Masukkan respon di sini | Masukkan respon di sini | Masukkan respon di sini | Masukkan respon di sini | Masukkan respon di sini | IDR | Masukkan respon di sini | Masukkan respon di sini | Masukkan respon di sini | Masukkan respon di sini | Masukkan respon di sini
IDR
IDR
IDR
IDR
IDR
IDR
IDR
IDR
IDR
IDR
IDR
IDR
IDR
IDR
IDR
Total Prices [Bagian ini otomatis terisi. Mohon agar tidak diubah]
[9 empty] | Mata Uang [Kolom ini akan terisi secara otomatis. Mohon jangan dirubah] | Harga [Kolom ini akan terisi secara otomatis. Mohon jangan dirubah]
[7 empty] | Subtotal 1: Total untuk Tahun 6 |  | IDR | 0
[7 empty] | Subtotal 2: Total untuk Tahun 7 |  | IDR | 0
| Waktu Respon terhadap Kerusakan |  | a. Remote (jam) | Maksimal 1 Jam |  |  | Subtotal 3: Total untuk Tahun 8 |  | IDR | 0
|  |  | b. On-site (jam) | Maksimal 48 Jam |  |  | Subtotal 4: Total untuk Tahun 9 |  | IDR | 0
| Waktu yang Dibutuhkan untuk Pengiriman / Pengganti Suku Cadang (jam) |  |  | ≥ 10 tahun |  |  | Subtotal 5: Total untuk Tahun 10 |  | IDR | 0
| Ketersediaan Unit Pengganti (Ya/Tidak) |  |  | Masukkan respon di sini |  |  | TOTAL HARGA: Subtotals 1 + 2 + 3 + 4 + 5 |  | IDR | 0
| Ketersediaan Teknisi Layanan (Ya/Tidak) |  |  | Ya
| Ketersediaan Layanan pada Hari Sabtu, Minggu, dan Hari Libur Nasional (Ya/Tidak) |  |  | Ya
Catatan untuk Bidder:
Layanan pasca masa garansi(Post Warranty-PW) yang tercantum dalam tabel ini harus mencakup layanan yang sama dengan layanan masa garansi, yang akan diberlakukan pada tahun ke-6 hingga ke-10. Meskipun harga penawaran layanan ini tidak mewakili kewajiban kontrak bagi Pembeli pada masa Post Warranty, penyedia tetap wajib menyediakan layanan operasi dan pemeliharaan perangkat, dan dapat dilakukan penyesuaian harga sebagaimana formula yang ditetapkan dalam kontrak.

## 1. Summary Cost
Ringkasan Harga
1. Harga Penawaran
Catatan untuk Peserta Lelang: Jumlah total dari daftar harga lainnya [(c) (d) (e) + (f)] dan Total Harga Penawaran [(a) dan (b)], untuk dicocokkan dengan jumlah dalam Surat Penawaran - Bagian Keuangan
MATA UANG (Sesuai DPP 18)
IDR
Kolom-kolom di bawah ini akan terisi otomatis. Mohon untuk tidak mengubah
(a) | Surat Penawaran - Bagian Keuangan (Termasuk Pajak) Perhitungan (sebagai referensi): (c) + (d) + (e) + (f) | 0
(b) | Surat Penawaran - Bagian Keuangan (Tidak Termasuk Pajak) Perhitungan (sebagai referensi): (a) - (g) - (h) - (i) | 0
(c) | Daftar Harga: Barang yang akan di Impor | 0
(d) | Daftar Harga: Barang yang sudah di Impor | 0
(e) | Daftar Harga: Barang Produksi Dalam Negeri | 0
(f) | Harga dan Jadwal Penyelesaian - Jasa Terkait | 0
2. Pengecualian pajak sesuai dengan DPP 20
Kolom-kolom di bawah ini akan terisi otomatis. Mohon untuk tidak mengubah
(g) | Jumlah pajak penjualan dan pajak serupa lainnya, yang akan dibayarkan atas barang jika kontrak diberikan kepada Penawar, yang termasuk dan dihargai dalam (e) | 0
(h) | Jumlah bea masuk dan pajak impor lainnya yang dikenakan atas Barang yang diimpor, pajak penjualan dan pajak serupa lainnya, termasuk dalam (c) | 0
(i) | Jumlah bea cukai dan pajak impor lainnya yang dikenakan atas Barang yang diimpor, pajak penjualan dan pajak serupa lainnya, termasuk dalam (d) | 0

## 2. PS-Goods to be Imported
Daftar Harga: Barang Yang Akan Di Impor
(Penawaran Kelompok C, barang yang akan diimpor) Mata uang sesuai dengan DPP 18 | [13 empty] | Tanggal :[Masukan Tanggal]
RFQ No: [Masukan Nomor RFQ]
Alternative No: [Masukan Nomor Alternative]
Nomor Halaman (dari total halaman):
1 | 2 | 3 | 4 | 5 |  | 6 |  | 7 |  | 8 |  | 9 |  | 10
No Baris | Deskripsi Barang | Negara Asal | Tanggal pengiriman sebagaimana ditentukan oleh Incoterms | Kuantitas dan nama fisik unit |  | Harga unit DDP setelah dikurangi bea masuk dan pajak impor, transportasi darat, asuransi, dan layanan lain yang diperlukan di Indonesia untuk mengirimkan barang ke tujuan akhir setiap baris |  | Bea masuk dan pajak impor yang harus dibayar per unit sebagaimana ditetapkan dalam DPP 17 (b) (ii) |  | *) Harga per item baris untuk transportasi darat, asuransi, dan layanan lain yang diperlukan di Indonesia untuk mengirimkan barang ke tujuan akhir, sebagaimana ditetapkan dalam DPP 17 (b) (iii) |  | Harga satuan DDP [masukkan harga ke tempat tujuan] sebagaimana ditetapkan dalam DPP 17 (b) (i) |  | Total Harga Per Baris (Kol 5x(6+7+8)]
[masukkan nomor dari item barang] | [masukkan nama dari barang] | [masukkan nama negara asal Barang] | [masukkan Tanggal Pengiriman yang dikutip] | [masukkan jumlah unit yang akan dipasok] | [masukkan nama unit fisik] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [masukkan harga unit DDP setelah dikurangi Bea Masuk dan Pajak Impor, transportasi darat, asuransi, dan jasa lain yang diperlukan di Indonesia untuk mengirimkan barang ke tujuan akhir per baris] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [Masukan Bea Masuk dan Pajak Impor yang harus dibayar per unit] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [masukkan harga per baris item untuk asuransi transportasi darat, dan layanan lain yang diperlukan di Negara Pembeli] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [masukkan harga unit DDP per unit] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | Harga unit DDP setelah dikurangi Bea Masuk dan Pajak Impor, transportasi darat, asuransi, dan jasa lain yang diperlukan di Negara Pembeli untuk mengirimkan barang ke tujuan akhir (Kolom 5x6). [Kolom ini akan otomatis terisi, mohon tidak mengunbah] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | Jumlah Bea Masuk dan Pajak Impor yang harus dibayar per-unit (Kolom 5*7). [Kolom ini akan otomatis terisi, mohon tidak mengunbah] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | Harga per baris item untuk asuransi transportasi darat, dan layanan lain yang diperlukan di Negara Pembeli (Kolom 5*8). [Kolom ini akan otomatis terisi, mohon tidak mengunbah]
1 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0
2 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0
3 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0
4 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0
5 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0
6 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0
7 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0
8 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0
9 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0
10 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0
11 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0
12 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0
13 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0
*) Rincian biaya diperlukan hanya untuk keperluan penerapan preferensi domestik, jika berlaku
Sub-Totals dan Jumlah Harga [Bagian ini akan otomatis terisi, harap tidak di edit]
[17 empty] | Mata Uang [Kolom ini akan otomatis terisi, mohon tidak mengunbah] | Jumlah Harga [Kolom ini akan otomatis terisi, mohon tidak mengunbah]
[11 empty] | Subtotal 1: Harga Unit DDP setelah dikurangi Bea Masuk dan Pajak Impor, transportasi darat, asuransi, dan jasa lain yang diperlukan (Jumlah dari Kolom P) | [5 empty] | IDR | 0
[11 empty] | Subtotal 2: Bea Masuk dan Pajak Impor (Jumlah dari Kolom R) | [5 empty] | IDR | 0
[11 empty] | Subtotal 3: Harga untuk Asuransi Transportasi Darat, dan layanan lain yang diperlukan di Negara Pembeli | [5 empty] | IDR | 0
Nama Penyedia [masukkan nama lengkap Penyedia] Tanda Tangan Penyedia [tanda tangan orang yang menandatangani Penawaran] Tanggal [Masukkan Tanggal] | [10 empty] | TOTAL PRICE: Subtotal 1 + Subtotal 2 + Subtotal 3 | [5 empty] | IDR | 0

## 3. PS- Goods already imported
Daftar Harga: Barang Yang Sudah Di Impor
(Penawaran Kelompok C, Barang yang sudah diimpor) Mata uang sesuai dengan DPP 18 | [17 empty] | Tanggal :[Masukan Tanggal]
RFQ No: [Masukan Nomor RFQ]
Alternative No: [Masukan Nomor Alternative]
Nomor Halaman (dari total halaman):
1 | 2 | 3 | 4 | 5 |  | 6 |  | 7 |  | 8 |  | 9 |  | 10 |  | 11 |  | 12
No Baris Item | Deskripsi Barang | Negara Asal | Tanggal Pengiriman sebagaimana ditentukan oleh Incoterms | Kuantitas dan nama unit fisik |  | Harga unit DDP setelah dikurangi bea masuk dan pajak impor, sebagaimana ditetapkan dalam DPP 17 (c) (iii) |  | Bea Masuk dan Pajak Impor yang dibayarkan per unit sebagaimana ditetapkan dalam DPP 17 (c)(ii), [harus didukung dengan dokumen] |  | Harga unit DDP termasuk Bea Masuk dan Pajak Impor yang dibayarkan, sebagaimana ditetapkan dalam DPP 17 (c) (i) |  | Harga setiap item setelah dikurangi bea masuk dan pajak impor yang dibayarkan, sebagaimana ditetapkan dalam DPP 17 (c)(i) (Kolom 5×6) |  | Harga satuan untuk transportasi darat, asuransi, dan layanan lain yang diperlukan di Negara Pembeli untuk mengirimkan barang ke tujuan akhir, sesuai dengan ketentuan sebagaimana ditetapkan dalam DPP 17 (c)(v) |  | Pajak penjualan dan pajak lainnya yang dibayarkan atau terutang setiap item jika Kontrak diberikan (sebagaimana ditetapkan dalam DPP 17 (c)(iv)) |  | Harga Total per Baris [Col. 9+5 x (10+7+11)]
[Masukan nomor] | [masukkan Nama Barang] | [masukkan Negara Asal Barang] | [masukkan Tanggal Pengiriman yang dikutip] | [masukkan jumlah unit yang akan dipasok] | [masukkan nama unit fisik] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [Masukkan harga unit DDP setelah dikurangi bea masuk dan pajak impor] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [Masukkan bea masuk dan pajak yang dibayarkan per satuan] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [Masukkan harga satuan DDP per satuan] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [Harga per item baris setelah dikurangi bea masuk dan pajak impor. Kolom ini akan terisi otomatis. Mohon untuk tidak mengubah.] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | Masukkan harga per unit untuk transportasi darat, asuransi, dan layanan lain yang diperlukan di Negara Pembeli] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | Masukkan pajak penjualan dan pajak lainnya yang harus dibayar per unit jika Kontrak diberikan | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | Harga per item baris, setelah dikurangi Bea Masuk dan Pajak Impor yang telah dibayarkan (Kolom 9). Kolom ini akan otomatis terupdate, mohon tidak mengedit | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | Harga per item baris untuk transportasi darat, asuransi, dan layanan lain yang diperlukan di Negara Pembeli (Kolom 5*10). Kolom ini akan otomatis terupdate, mohon tidak mengedit | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | Masukkan harga per item baris untuk Bea Masuk dan Pajak Impor yang telah dibayarkan (Kolom 5*7). Kolom ini akan otomatis terupdate, mohon tidak mengedit | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | Masukkan harga per item baris untuk Pajak Penjualan dan pajak lainnya yang dibayarkan atau harus dibayarkan per item jika Kontrak diberikan (Kolom 5*11). Kolom ini akan otomatis terupdate, mohon tidak mengedit
1 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0 | IDR | 0
2 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0 | IDR | 0
3 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0 | IDR | 0
4 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0 | IDR | 0
5 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0 | IDR | 0
6 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0 | IDR | 0
7 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0 | IDR | 0
8 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0 | IDR | 0
9 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0 | IDR | 0
10 | [5 empty] | IDR |  | IDR |  | IDR |  | IDR | 0 | IDR |  | IDR |  | IDR | 0 | IDR | 0 | IDR | 0 | IDR | 0
Sub-Totals dan Jumlah Harga [Bagian ini akan otomatis terisi, harap tidak di edit]
[23 empty] | Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit | Harga Kolom ini akan otomatis terupdate, mohon tidak mengedit
[17 empty] | Subtotal 1: Harga Unit DDP, setelah dikurangi Bea Masuk dan Pajak Impor yang telah dibayarkan (Jumlah dari kolom T) | [5 empty] | IDR | 0
[17 empty] | Subtotal 2: Transportasi darat, Asuransi, dan layanan lain yang diperlukan di Negara Pembeli (Jumlah dari kolom V) | [5 empty] | IDR | 0
Nama Penyedia [masukkan nama lengkap Penyedia] Tanda Tangan Penyedia [tanda tangan orang yang menandatangani Penawaran] Tanggal [Masukkan Tanggal] | [16 empty] | Subtotal 3: Bea Masuk dan Pajak Impor yang telah dibayarkan (Jumlah dari kolom X) | [5 empty] | IDR | 0
[17 empty] | Subtotal 4: Pajak Penjualan dan pajak lainnya yang dibayarkan atau harus dibayarkan per item jika Kontrak diberikan (Jumlah dari Kolom Z) | [5 empty] | IDR | 0
[17 empty] | Subtotal 5: Jumlah bea masuk dan pajak impor lainnya yang dikenakan pada Barang impor, pajak penjualan, dan pajak serupa lainnya (Subtotal 3 + Subtotal 4) | [5 empty] | IDR | 0
[17 empty] | HARGA TOTAL: Subtotal 1 + Subtotal 2 + Subtotal 3 + Subtotal 4 | [5 empty] | IDR | 0
* [Untuk Barang yang sudah di Impor, harga penawaran harus dipisahkan dari nilai impor Barang tersebut yang akan dikenakan cukai dan harus memasukkan nilai pemotongan atau penambahan dari agen atau perwakilan di Indonesia dan seluruh biaya lokal lainnya selain biaya cukai dan pajak impor , yang telah dan/atau harus dibayarkan oleh Purchaser (Kementerian Kesehatan. Untuk lebih jelasnya, Peserta diminta untuk memasukan harga penawaran termasuk biaya cukai dan pajak impor, dan sebagai tambahan juga memberikan informasi biaya cukai impor dan harga Barang yang berbeda dengan nilai tersebut diatas.]

## 4. PS-Goods - Purchaser's Cntry
Price Schedule: Goods Manufactured in the Purchaser’s Country
Negara Pembeli | [13 empty] | Tanggal :[Masukan Tanggal]
______________________ | [13 empty] | RFQ No: [Masukan Nomor RFQ]
Alternative No: [Masukan Nomor Alternative]
Nomor Halaman (dari total halaman):
1 | 2 | 3 | 4 |  | 5 |  | 6 |  | 7 |  | 8 |  |  | 9 |  | 10
Nomor Item Baris | Deskripsi Barang | Tanggal pengiriman sebagaimana ditentukan oleh Incoterms | Kuantitas dan nama unit fisik |  | Harga unit DDP termasuk harga untuk transportasi darat, asuransi, dan layanan lokal lainnya yang diperlukan untuk mengirimkan Barang ke tujuan akhir |  | Pajak penjualan dan pajak lainnya yang harus dibayarkan per unit jika Kontrak diberikan (sebagaimana ditetapkan dalam DPP 17 (a)(ii) |  | Harga satuan DDP setelah dikurangi Pajak Penjualan dan pajak lainnya yang harus dibayarkan per unit |  | Biaya tenaga kerja lokal, bahan baku, dan komponen yang berasal dari negara asal di Indonesia sebagai persentase dari Kol. 5 |  |  | Total pajak penjualan dan pajak lainnya (Kolom 4x6) |  | Total Harga setiap item per baris (Kolom 4x7)
[Masukan Nomor) | [Masukan nama Barang] | [masukkan Tanggal Pengiriman yang dikutip] | [masukkan jumlah unit yang akan dipasok] | [masukkan nama unit fisik] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [masukkan harga satuan DDP termasuk harga untuk transportasi darat, asuransi, dan layanan lokal lainnya yang diperlukan untuk mengirimkan Barang ke tujuan akhir] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [masukkan pajak penjualan dan pajak lainnya yang harus dibayarkan setiap item per baris jika Kontrak diberikan] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [masukkan harga satuan setelah dikurangi pajak penjualan dan pajak lain yang harus dibayar (Col. 5 - Col. 6). Kolom ini akan otomatis terupdate, mohon tidak mengedit | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [masukkan %] | [Masukkan biaya tenaga kerja lokal, bahan baku, dan komponen dari dalam negara tempat Pembelian sebagai persentase dari harga DDP setiap item per baris. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [Total pajak penjualan dan pajak lainnya. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [Total harga setiap item per baris. Kolom ini akan otomatis terupdate, mohon tidak mengedit]
1 |  |  |  |  | IDR |  | IDR |  | IDR | 0 | IDR |  | 0 | IDR | 0 | IDR | 0
2 |  |  |  |  | IDR |  | IDR |  | IDR | 0 | IDR |  | 0 | IDR | 0 | IDR | 0
3 |  |  |  |  | IDR |  | IDR |  | IDR | 0 | IDR |  | 0 | IDR | 0 | IDR | 0
4 |  |  |  |  | IDR |  | IDR |  | IDR | 0 | IDR |  | 0 | IDR | 0 | IDR | 0
5 |  |  |  |  | IDR |  | IDR |  | IDR | 0 | IDR |  | 0 | IDR | 0 | IDR | 0
6 |  |  |  |  | IDR |  | IDR |  | IDR | 0 | IDR |  | 0 | IDR | 0 | IDR | 0
7 |  |  |  |  | IDR |  | IDR |  | IDR | 0 | IDR |  | 0 | IDR | 0 | IDR | 0
8 |  |  |  |  | IDR |  | IDR |  | IDR | 0 | IDR |  | 0 | IDR | 0 | IDR | 0
9 |  |  |  |  | IDR |  | IDR |  | IDR | 0 | IDR |  | 0 | IDR | 0 | IDR | 0
10 |  |  |  |  | IDR |  | IDR |  | IDR | 0 | IDR |  | 0 | IDR | 0 | IDR | 0
Sub-Totals dan Jumlah Harga [Bagian ini akan otomatis terisi, harap tidak di edit]
[16 empty] | Mata Uang.Kolom ini akan otomatis terupdate, mohon tidak mengedit | Harga [Kolom ini akan otomatis terupdate, mohon tidak mengedit]
Nama Penyedia [masukkan nama lengkap Penyedia] Tanda Tangan Penyedia [tanda tangan orang yang menandatangani Penawaran] Tanggal [Masukkan Tanggal] | [9 empty] | Subtotal 1: Total Pajak Penjualan dan pajak lainnya yang dibayarkan atau harus dibayarkan jika Kontrak diberikan (Jumlah dari Kolom 9) | [5 empty] | IDR | 0
[10 empty] | Subtotal 2: Total Harga tidak termasuk pajak penjualan dan pajak lainnya (Jumlah dari Kolom 10) | [5 empty] | IDR | 0
[10 empty] | Total Harga: Subtotal 1 + Subtotal 2 | [5 empty] | IDR | 0

## 5. PS-Related Services
Daftar Harga - Layanan Terkait
Mata uang sebagaimana ditetapkan dalam DPP 18 | [7 empty] | Tanggal :[Masukan Tanggal]
RFQ No: [Masukan Nomor RFQ]
Alternative No: [Masukan Nomor Alternative]
Nomor Halaman (dari total halaman):
1 | 2 | 3 | 4 | 5 |  | 6 |  | 7
Nomor Jasa | Deskripsi Layanan (tidak termasuk transportasi darat dan layanan lainnya yang diperlukan di Negara Pembeli untuk mengirimkan barang ke tujuan akhirnya). | Negara Asal | Tanggal Pengiriman di tempat tujuan akhir | Kuantitas dan satuan fisik |  | Harga Unit |  | Total Harga per Layanan (Kol. 5*6)
Masukan Nomor Jasa | [Masukkan nama Layanan yang sesuai dengan Daftar Layanan Terkait dan Jadwal Penyelesaian di Bagian VII] | [Masukan Negara Asal Layanan] | [masukkan tanggal pengiriman di tempat tujuan akhir per Layanan] | [masukkan jumlah unit yang akan dipasok] | [masukkan nama fisik unit] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [Masukan harga unit] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | [Total Harga per Layanan. Kolom ini akan otomatis terupdate, mohon tidak mengedit]
1 | Instalasi, Uji Coba Operasional dan komisioning |  |  |  |  | IDR |  | IDR | 0
2 | Pelatihan |  |  |  |  | IDR |  | IDR | 0
3 | Garansi untuk Pemeliharaan dan Perbaikan Tahun 1 hingga Tahun 5 |  |  |  |  | IDR |  | IDR | 0
4 | Proposed Consumables: Total biaya starter kit Penyedia diminta memberikan informasi ini dalam "Proposed Consumables & Price" sheet. Cells terkait untuk "Harga Satuan" (Kol. 6) dan "Total Harga per Layanan" (Kol. 7) akan otomatis terupdate. Mohon tidak mengedit. |  |  |  |  | IDR | This column will be auto-filled. Kindly do not edit | IDR | 0
| Proposed Consumables: Total proyeksi biaya tambahan selama 5 tahun Penyedia diminta memberikan informasi ini dalam "Proposed Consumables & Price" sheet. Cells terkait untuk "Harga Satuan" (Kol. 6) dan "Total Harga per Layanan" (Kol. 7) akan otomatis terupdate. Mohon tidak mengedit. |  |  |  |  | IDR | This column will be auto-filled. Kindly do not edit | IDR | 0
Harga Total [Bagian ini akan otomatis terupdate, mohon tidak mengedit]
[8 empty] | [Mata Uang. Kolom ini akan otomatis terupdate, mohon tidak mengedit] | Harga [Kolom ini akan otomatis terupdate, mohon tidak mengedit]
Catatan untuk Peserta: Harga bahan habis pakai eksklusif bersifat tetap dan mengikat selama lima tahun (termasuk pengiriman) dimulai dari tanggal Pengujian Penerimaan & Uji coba yang berhasil. Penyediaan starter kit termasuk dalam harga penawaran. Total perkiraan biaya bahan habis pakai selama periode 5 tahun sebagaimana dikutip oleh peserta dalam Daftar Harga Bahan Habis Pakai yang Diusulkan dan Harga (5 tahun) akan ditambahkan ke Harga Penawaran untuk tujuan evaluasi. Biaya ini tidak akan dimasukkan dalam Harga Penawaran atau Jumlah Kontrak. | [5 empty] | HARGA TOTAL (Hasil Penjumlahan dari Kol. 9, tidak termasuk proyeksi biaya tambahan selama 5 tahun / "additional sets of proposed consumables") |  | IDR | 0
Nama Penyedia [masukkan nama lengkap Penyedia] Tanda Tangan Penyedia [tanda tangan orang yang menandatangani Penawaran] Tanggal [Masukkan Tanggal]